**ON/OFF** Key-Press to turn meter on.

**Sample/End** Key-Press and release to start the sample, then press the **End** key to stop the sample.

**Stats** Key-Press to view the minimum, maximum and average of the last sample taken.



 Battery Indicator

 CO2 Level in PPM

 **Sample/End** Key **Calibration** Key

  **Stats** Key **ON/OFF** Key

 **Calibration** Key-Press to begin field calibration. First connect zero gas to the instrument and open the regulator. Then use the **↑, ↓** , and **ENTER** keys to select **ZERO**. The instrument will automatically zero itself after one minute. (The three buttons are redefined to the label on the LCD above each button)

After zeroing the sensor, connect span gas to the instrument and open the regulator. Then use the **↑, ↓** , and **ENTER** keys to select **SPAN.** The instrument will take one minute to provide a reading. Next use the **↑, ↓** , and **ENTER** keys to adjust the concentration displayed on the instrument to match the concentration on the gas cylinder.