DOCUMENT TYPE:		Standard Operating Procedure

TITLE:					Flow Check on BAM 1020 (Audit)

PREPARED BY:                          									7/03/13

Tribal Air Program					                       				       	(Date)

REVIEWED BY:

Environmental Technician		                               						(Date)

APPROVED BY:

Environmental Department Manager							         (Date)

Table of Contents

- Scope of Application……………………………………………………………………………………………….1

- Summary of Method………………………………………………………………………………………….…….1

- Interferences/Comments……………………………………………………………………………………………1

- Supplies………………………………………………………………………………………………….………….2

- Procedures Setups………………………………………………………………………..……………………….2-3

- Troubleshooting…………………………………………………………………………….……………………….3

# Scope of Application

## Quality Control (QC) is the overall system of technical activities that measures the attributes and performance of a process, item, or service against defined standards to verify that they meet the stated requirements established by the tribe. Day-to-day quality control is implemented through the use of various checks or instruments that are used for comparison.

# Summary of Method

## Place BAM 1020 in “maintenance mode”

## Correct readings

## Measure volumetric flow

## Reassemble BAM 1020

# Interferences/Comments

## The volumetric flow check requires the use of the BX-307 Flow Calibrator.  Consult the SOP titled “Operating the BX-307”.

  - Using the BAM-1020 Keypad/Display: This main menu screen is the starting point for all functions of the BAM and can be accessed from any screen by selecting EXIT.

    - Soft keys: Four white buttons directly beneath the display changes in response to a menu option above.

    - Function keys: Used for entering password F1,F2,F3,F4

    - Arrow keys: Used to navigate up, down, left and right

  - A flow rate verification check of the sampler flow rate is recommended every month.  Control charts presenting flow rate verification check data (indicated vs. measured) should be maintained.

  - The below procedures are for flow “check” or “audit” in which no BAM calibrations are to be changed.

  - Compare the BAM column reading to your standard device for each parameter, and record the results.  No calibrations are altered if the CAL or DEFAULT keys are not pressed.

  - If calibration is required see SOP titled “Flow Calibration on BAM 1020”

# Supplies

## BX-307 Flow Calibrator (located in black case the workstation’s cabinet)

  - BAM-1020 Calibration Sheet (located on clipboard on west wall)

# Procedure

## Install BX-307Flow Calibrator

## On the shelter’s roof, turn on the BX-307 Flow Calibrator (standard-STD) and set it aside to allow the unit to self stabilize (this can be done before the leak check is performed).

## Remove the PM10 head (or BX-302 leak check valve, if performing after leak check) and the VSCC from inlet and install the BX-307 Flow Calibrator, wrapping the box’s cord around bi-pod’s leg as to keep the box from dangling freely.

## Obtain Temperature and Barometric Pressure readings

## Take note of the BX-307 temperature and barometric pressure readings and enter these two values on the BAM-1020 Calibration Sheet under “Ref. Temperature” and “Ref. Pressure”.

## Compare Temperature and Barometric Reading into BAM-1020 STD field

## At the BAM’s main menu screen press the TEST soft key then select FLOW.

## Enter password: F1, F2, F3, F4

## The “Multipoint Flow Calibration” screen will appear. Compare the BAM’s Ambient Temperature (AT) and Barometric Pressure (BP) to the values obtained from standard device, BX-307 in step 5.2.1

## If AT is within ±2°C select NEXT soft key.

## If BP is within ±10 mmHg, select NEXT soft key.

## Verify Flow 1,2, and 3

## Flow 1: On the BAM press the NEXT key to select the first flow point of 15.0 L/min.  The BAM-1020 pump will turn on automatically.  Allow the unit to regulate the flow until the BAM reading stabilizes at the target flow rate of 15.0 L/min.

## Return to the shelter’s roof and allow the BX-307 to stabilize.  Once stabilized, take note of the volumetric flow rate value (Qa).

## Return to the BAM-1020 and compare the standard device, BX-307 (Qa) value to the BAM field for FLOW 1.  These are simple checks and no calibrations are required.

## Flow 2: Press NEXT key to select the second flow point of 18.4 L/min.  Allow the flow to stabilize again.

## Take note of the new volumetric flow rate value (Qa) on the BX-307at the shelter’s roof.

## Compare this value to the BAM field for FLOW 2 and press the NEXT key.

## Flow 3: Press the NEXT key to select the third flow point of 16.7 L/min.  Allow the flow to stabilize then write the value from the BAM column for FLOW 3 onto the Calibration sheet under “BAM Flow”.

## At the shelter’s roof, check the volumetric flow rate value (Qa) on the BX-307 and enter this value onto the Calibration sheet under “Ref. Flow”.

## If the percentage difference is less than 2%, select the NEXT soft key.

## Remove BX307 and replace both PM heads.

## Complete BAM Calibration Checklist.

# Troubleshooting

## Refer to the BAM 1020 Particulate Monitor Operation Manual