DOCUMENT TYPE:		Standard Operating Procedure

TITLE:					BAM-1020 Inlet Tube Cleaning

PREPARED BY:                     						                        08/02/2011

Tribal Air Program					                            			  	(Date)

REVIEWED BY:

Environmental Supervisor                                  						(Date)

APPROVED BY:

Environmental Department Manager							         (Date)

Table of Contents

- Scope of Application……………………………………………………………………………………………….1

- Summary of Method………………………………………………………………………………………….…….1

- Interferences/Comments……………………………………………………………………………………………2

- Supplies………………………………………………………………………………………………….………….2

- Procedures Setups………………………………………………………………………..………………………….2

- Troubleshooting…………………………………………………………………………….……………………….2

# Scope of Application

## Gather ambient air data, which have been linked to smog, ozone depletion, acid rain and global warming. Maintaining the equipment that records this data is a vital aspect of providing consistent, current and valid data.

# Summary of Method

## Disassemble attachments to the inlet tube

## Remove inlet tube from BAM-1020

## Clean inlet tube

## Re-attach inlet tube to BAM-1020

## Re-assemble attachments

# Interferences/Comments

## The best time to perform this cleaning procedure is following the BAM leak and flow check as it eliminates some of the BAM down time.

## All supplies can be found in work station cabinet or desk drawer

# Supplies

## Kim-wipe

## Cloth/rag

## Cotton swab

## Red plastic inlet cap

## Isopropyl alcohol

# Procedures

## Place BAM into maintenance mode by pressing SET UP then CALIBRATE.

## Remove the heater cover from over the heater/inlet tube.

## Loosen the 2 set screws on the top of the BAM cabinet using the proper Allen key.

## From the roof of the shelter, remove the PM10 and SCC assembly by pulling straight off.

## Loosen the temperature/pressure assembly set screws then pull straight off/over the inlet tube.

## Using a screwdriver, loosen the clamp holding the bipod to the inlet tube.

## By hand, loosen the plastic flange just above the roof where the inlet tube meets the roof.

## Inside the shelter, loosen the two set screws of the heater assembly and slide it down the tube.

## Gently push the tube up far enough to clear it of the heater assembly then move the heater out of the way.

## Place red plastic cap over inlet receptive area to keep any debris from entering the BAM.

## From the roof of the shelter, gently pull inlet tube up and out.

## Clean with cloth running inside from one end to the other (Caution: avoid using any abrasive materials).

## After inlet tube is clean and dry, reassemble and re-install, reversing the above steps.

# Troubleshooting

Refer to the BAM-1020 9800 Manual Rev. G in m:\_103\Task 1_Air Monitoring\AMS_Equipment folder.