## BKGD Audit
Morongo Tribal Air Program
BAM1022 Background (BKGD) Audit
| Operator: |  |  | Location: |  |  |  | Sampler: | MetOne BAM1022
[8 empty] | Serial #: | C19683
| Pre-BKGD Value: | [6 empty] | Factory BKGD: | -0.0003
|  | Start | End |  |  | Leak Valve: | MetOne BX-305
Date: | [5 empty] | Zero Filter Assembly: | MetOne BX-302
Real Time: | [5 empty] | Flow Meter: | BGI DeltaCal/BX-307 Serial #: 405 Certify Date: 9/5/2023
Logger Time:
Leak Check
|  | BAM1022 |  | Reference |  | Acceptance |  | Difference |  | PASS/FAIL
| Leak Check: |  |  | Flow Control Valve (closed) |  | (<1.0 lpm) |  |  |  | PASS
|  | Deg. C |  | Deg. C
| Ambient Temp: |  |  |  |  | ± 2.1 Deg. C |  | 0 |  | PASS
|  | mmHg |  | mmHg
| Pressure: |  |  |  |  | ± 10.1 mmHG |  | 0 |  | PASS
|  | L/min |  | L/min
| Flow Rate: |  |  |  |  | ± 2.1% |  | #DIV/0! |  | #DIV/0!
| Design Flow: | 16.7 |  |  |  | ± 5.1% |  | -100 |  | FAIL
Pre-Verification
Verify the following: | [9 empty] | YES/NO
-1022 recommended maintenance performed periodically and prior to zero test?
-1022 filter tape spots do not have pin holes or PM present?
-Background value set to 0 before beginning BKGD Audit?
Post-Verification
Verify the following: | [9 empty] | YES/NO
-Raw data inputted into "Raw Data Input" tab?
-72 or more consecutive data points? |  |  |  |  | Counted?: | 72 |  |  |  | YES
-Investigated proper operations during background determination duration in "Investigate" tab?
-Checked "Zero Calculation" tab?
-Removed outliers on "Theoretical Zero" tab?
-Status on hourly data is 0? (0 means no problems) | [9 empty] | YES
Background
| Last BKGD Value: | 0 |  | New BKGD Value: | -0.00082361 |  |  |  | ΔBKGD Value: | 0.00082361
Notes/Comments:

## Raw Data Input
Time | ConcRT(ug/m3) | ConcHR(ug/m3) | ConcS(ug/m3) | Flow(lpm) | WS(m/s) | WD(Deg) | AT(C) | RH(%) | BP(mmHg) | FT(C) | FRH(%) | Status
2023-08-19 09:00 | -1.7 | 1 | 1 | 16.62 | 0.3 | 0 | 23.7 | 52 | 700 | 31.6 | 27 | 0
2023-08-19 10:00 | 3.9 | -0.5 | -0.6 | 16.61 | 0.3 | 1 | 23.8 | 51 | 700.2 | 31.6 | 27 | 0
2023-08-19 11:00 | -2.1 | 0 | 0 | 16.61 | 0.3 | 1 | 23.6 | 52 | 700.3 | 31.6 | 27 | 0
2023-08-19 12:00 | -2.8 | 1.1 | 1.2 | 16.61 | 0.3 | 1 | 23.6 | 51 | 699.8 | 31.6 | 26 | 0
2023-08-19 13:00 | -4 | 1.6 | 1.7 | 16.61 | 0.3 | 1 | 23.7 | 52 | 699.3 | 31.6 | 27 | 0
2023-08-19 14:00 | 3.7 | 3.1 | 3.4 | 16.61 | 0.3 | 1 | 23.5 | 50 | 698.8 | 31.7 | 26 | 0
2023-08-19 15:00 | -0.3 | -0.1 | -0.1 | 16.61 | 0.3 | 0 | 23.7 | 51 | 698.5 | 31.6 | 27 | 0
2023-08-19 16:00 | -1.4 | 0.2 | 0.3 | 16.61 | 0.3 | 0 | 23.7 | 50 | 698.3 | 31.7 | 26 | 0
2023-08-19 17:00 | -1.9 | -0.6 | -0.7 | 16.6 | 0.3 | 1 | 23.9 | 50 | 698.3 | 31.8 | 26 | 0
2023-08-19 18:00 | -0.6 | 0.9 | 1 | 16.61 | 0.3 | 1 | 23.9 | 50 | 698.2 | 31.8 | 26 | 0
2023-08-19 19:00 | -1.4 | 0.4 | 0.4 | 16.61 | 0.3 | 0 | 24.2 | 51 | 698.2 | 32 | 27 | 0
2023-08-19 20:00 | 0.8 | -2.1 | -2.3 | 16.62 | 0.3 | 0 | 24.3 | 51 | 698.2 | 32.1 | 27 | 0
2023-08-19 21:00 | -0.6 | -1.9 | -2 | 16.61 | 0.3 | 0 | 24.2 | 50 | 698.5 | 32.1 | 26 | 0
2023-08-19 22:00 | -0.4 | 0.3 | 0.3 | 16.61 | 0.3 | 0 | 24.4 | 50 | 698.6 | 32.3 | 26 | 0
2023-08-19 23:00 | 0 | 3.7 | 4.1 | 16.61 | 0.3 | 0 | 24.5 | 50 | 698.2 | 32.3 | 26 | 0
2023-08-20 | 1.9 | 1.2 | 1.3 | 16.6 | 0.3 | 0 | 24.6 | 51 | 698 | 32.3 | 27 | 0
2023-08-20 01:00 | -0.9 | -0.9 | -0.9 | 16.61 | 0.3 | 1 | 24.6 | 52 | 697.7 | 32.3 | 27 | 0
2023-08-20 02:00 | -2.5 | -1 | -1.1 | 16.61 | 0.3 | 1 | 24.6 | 52 | 697.6 | 32.5 | 28 | 0
2023-08-20 03:00 | -1.6 | 0.5 | 0.5 | 16.61 | 0.3 | 1 | 24.6 | 51 | 697.6 | 32.4 | 27 | 0
2023-08-20 04:00 | 3.2 | 0.3 | 0.3 | 16.62 | 0.3 | 0 | 24.7 | 51 | 696.7 | 32.4 | 27 | 0
2023-08-20 05:00 | -0.7 | 0.7 | 0.8 | 16.6 | 0.3 | 0 | 24.8 | 51 | 696.3 | 32.5 | 27 | 0
2023-08-20 06:00 | -0.7 | 0.6 | 0.7 | 16.6 | 0.3 | 0 | 24.5 | 52 | 696 | 32.4 | 27 | 0
2023-08-20 07:00 | 3.8 | 2.2 | 2.4 | 16.62 | 0.3 | 0 | 24.7 | 51 | 695.7 | 32.3 | 27 | 0
2023-08-20 08:00 | -1.8 | 0 | 0 | 16.61 | 0.3 | 1 | 24.8 | 52 | 695.4 | 32.5 | 27 | 0
2023-08-20 09:00 | 4.6 | 3.3 | 3.6 | 16.62 | 0.3 | 0 | 24.4 | 54 | 695 | 32.2 | 28 | 0
2023-08-20 10:00 | 1.3 | 0.4 | 0.5 | 16.62 | 0.3 | 1 | 24.6 | 53 | 694.5 | 32.3 | 28 | 0
2023-08-20 11:00 | 5.9 | 6 | 6.5 | 16.61 | 0.3 | 1 | 24.2 | 54 | 694.6 | 32.1 | 28 | 0
2023-08-20 12:00 | 6 | 0.6 | 0.6 | 16.61 | 0.3 | 0 | 24.5 | 53 | 693.9 | 32.2 | 28 | 0
2023-08-20 13:00 | 2.4 | 1 | 1.1 | 16.61 | 0.3 | 0 | 24.4 | 55 | 692.3 | 32.1 | 28 | 0
2023-08-20 14:00 | 1 | 1.7 | 1.9 | 16.61 | 0.3 | 1 | 24.4 | 54 | 691.9 | 32.1 | 28 | 0
2023-08-20 15:00 | 2.4 | -0.8 | -0.8 | 16.62 | 0.3 | 0 | 24.2 | 58 | 691.1 | 31.9 | 30 | 0
2023-08-20 16:00 | 2.1 | 4.4 | 4.8 | 16.62 | 0.3 | 0 | 24.3 | 57 | 690.6 | 32 | 29 | 0
2023-08-20 17:00 | 2.6 | 1.3 | 1.4 | 16.62 | 0.3 | 0 | 24.2 | 58 | 691 | 31.8 | 30 | 0
2023-08-20 18:00 | 2 | 0.7 | 0.8 | 16.61 | 0.3 | 1 | 24.7 | 56 | 692.2 | 32.2 | 29 | 0
2023-08-20 19:00 | -2.1 | 0 | 0 | 16.62 | 0.3 | 1 | 24.5 | 55 | 693.1 | 32.2 | 29 | 0
2023-08-20 20:00 | 2.1 | -0.4 | -0.5 | 16.62 | 0.3 | 1 | 24.7 | 55 | 694.3 | 32.3 | 29 | 0
2023-08-20 21:00 | 0.8 | 0.6 | 0.7 | 16.61 | 0.3 | 1 | 24.7 | 54 | 696.3 | 32.4 | 28 | 0
2023-08-20 22:00 | -1.6 | 2.6 | 2.8 | 16.62 | 0.3 | 0 | 24.8 | 54 | 697.9 | 32.2 | 29 | 0
2023-08-20 23:00 | 1.6 | 1.8 | 1.9 | 16.62 | 0.3 | 1 | 24.9 | 53 | 698.7 | 32.3 | 28 | 0
2023-08-21 | 0.6 | 0.3 | 0.3 | 16.63 | 0.3 | 0 | 25.2 | 51 | 699.9 | 32.3 | 28 | 0
2023-08-21 01:00 | -1.6 | 2.7 | 2.9 | 16.61 | 0.3 | 0 | 25.3 | 49 | 700.6 | 32.4 | 26 | 0
2023-08-21 02:00 | 0.3 | -1 | -1.1 | 16.61 | 0.3 | 0 | 24.9 | 48 | 700.7 | 32.1 | 26 | 0
2023-08-21 03:00 | -3.6 | 0.5 | 0.6 | 16.59 | 0.3 | 0 | 25.5 | 46 | 700.9 | 32.4 | 25 | 0
2023-08-21 04:00 | -1.4 | -0.4 | -0.4 | 16.62 | 0.3 | 1 | 25.6 | 45 | 701.3 | 32.4 | 25 | 0
2023-08-21 05:00 | 1.2 | 2.2 | 2.3 | 16.61 | 0.3 | 0 | 25.6 | 45 | 701.5 | 32.4 | 25 | 0
2023-08-21 06:00 | -4.7 | 0.4 | 0.4 | 16.59 | 0.3 | 0 | 25.3 | 45 | 701.9 | 32.3 | 25 | 0
2023-08-21 07:00 | 3.2 | 2.7 | 2.9 | 16.62 | 0.3 | 1 | 24.7 | 48 | 702.4 | 31.9 | 26 | 0
2023-08-21 08:00 | -2.6 | 1 | 1.1 | 16.62 | 0.3 | 1 | 24.8 | 48 | 702.5 | 32.1 | 26 | 0
2023-08-21 09:00 | -0.3 | 0.2 | 0.2 | 16.61 | 0.3 | 0 | 24.2 | 48 | 702.9 | 32 | 25 | 0
2023-08-21 10:00 | 1.1 | 0 | 0 | 16.62 | 0.3 | 0 | 24.2 | 48 | 702.8 | 32 | 25 | 0
2023-08-21 11:00 | 2 | 0 | 0 | 16.61 | 0.3 | 0 | 23.7 | 50 | 702.6 | 31.6 | 26 | 0
2023-08-21 12:00 | -1.6 | 0.2 | 0.2 | 16.61 | 0.3 | 0 | 23.6 | 51 | 702.3 | 31.4 | 27 | 0
2023-08-21 13:00 | 1.8 | 0.5 | 0.6 | 16.61 | 0.3 | 0 | 23.7 | 51 | 701.9 | 31.4 | 27 | 0
2023-08-21 14:00 | -2.2 | -1.7 | -1.8 | 16.61 | 0.3 | 0 | 23.9 | 51 | 701.6 | 31.6 | 27 | 0
2023-08-21 15:00 | -4.1 | 1.6 | 1.8 | 16.6 | 0.3 | 0 | 24 | 51 | 701.3 | 31.7 | 27 | 0
2023-08-21 16:00 | 2.1 | 2.9 | 3.1 | 16.6 | 0.3 | 0 | 24.2 | 50 | 701.2 | 31.7 | 26 | 0
2023-08-21 17:00 | 3.9 | 2.7 | 2.9 | 16.6 | 0.3 | 0 | 24.2 | 50 | 700.9 | 31.8 | 26 | 0
2023-08-21 18:00 | 3.8 | 2.3 | 2.5 | 16.6 | 0.3 | 1 | 24.3 | 50 | 700.8 | 31.9 | 26 | 0
2023-08-21 19:00 | -2.7 | 0 | 0 | 16.61 | 0.3 | 0 | 24.4 | 49 | 701 | 32.1 | 26 | 0
2023-08-21 20:00 | -2 | 0.4 | 0.5 | 16.6 | 0.3 | 1 | 24.5 | 50 | 701.5 | 32.2 | 26 | 0
2023-08-21 21:00 | 2.1 | -1.6 | -1.7 | 16.6 | 0.3 | 1 | 24.6 | 50 | 701.9 | 32.2 | 26 | 0
2023-08-21 22:00 | 3.3 | 0 | 0 | 16.61 | 0.3 | 0 | 25 | 50 | 702 | 32.4 | 26 | 0
2023-08-21 23:00 | 1.1 | 0 | 0 | 16.62 | 0.3 | 0 | 25 | 49 | 702.1 | 32.4 | 26 | 0
2023-08-22 | -3.2 | -1.2 | -1.4 | 16.61 | 0.3 | 1 | 25.1 | 48 | 702.1 | 32.6 | 25 | 0
2023-08-22 01:00 | -0.7 | 3.7 | 4 | 16.61 | 0.3 | 1 | 25.3 | 48 | 701.8 | 32.6 | 26 | 0
2023-08-22 02:00 | 0 | 2.5 | 2.7 | 16.6 | 0.3 | 0 | 25.3 | 47 | 701.4 | 32.7 | 25 | 0
2023-08-22 03:00 | -0.1 | -0.7 | -0.8 | 16.63 | 0.3 | 0 | 25.4 | 47 | 701.2 | 32.6 | 25 | 0
2023-08-22 04:00 | -2.3 | 1 | 1.1 | 16.6 | 0.3 | 0 | 25.9 | 45 | 700.9 | 32.9 | 25 | 0
2023-08-22 05:00 | 1.9 | 2.3 | 2.4 | 16.63 | 0.3 | 0 | 25.3 | 46 | 700.9 | 32.6 | 25 | 0
2023-08-22 06:00 | 0.9 | 1.4 | 1.6 | 16.59 | 0.3 | 1 | 26.1 | 45 | 701.2 | 32.8 | 24 | 0
2023-08-22 07:00 | 1.9 | 0.8 | 0.9 | 16.62 | 0.3 | 0 | 25.4 | 46 | 701.5 | 32.7 | 24 | 0
2023-08-22 08:00 | -1.3 | -0.3 | -0.3 | 16.61 | 0.3 | 0 | 24.1 | 47 | 701.6 | 32.2 | 24 | 0

## Investigate
Time | ConcHR(ug/m3) | Flow(lpm) | AT(C) | FT(C) | RH(%) | FRH(%) | ΔAT | ΔFT | ΔRH | ΔFRH | AT<FT | RH>FRH | Investigate proper operations during background determination duration: |  |  | PASS/FAIL
2023-08-19 09:00 | 1 | 16.62 | 23.7 | 31.6 | 52 | 27 |  |  |  |  | PASS | PASS | BKGD value deviation < ± 2 µg/m3?: | -σ: | 0.0015123 | PASS
2023-08-19 10:00 | -0.5 | 16.61 | 23.8 | 31.6 | 51 | 27 | 0.1 | 0 | -1 | 0 | PASS | PASS | Filter temperature (FT) > Ambient Temperature (AT) ≈ 5-10 ˚C |  |  | PASS
2023-08-19 11:00 | 0 | 16.61 | 23.6 | 31.6 | 52 | 27 | -0.2 | 0 | 1 | 0 | PASS | PASS
2023-08-19 12:00 | 1.1 | 16.61 | 23.6 | 31.6 | 51 | 26 | 0 | 0 | -1 | -1 | PASS | PASS
2023-08-19 13:00 | 1.6 | 16.61 | 23.7 | 31.6 | 52 | 27 | 0.1 | 0 | 1 | 1 | PASS | PASS
2023-08-19 14:00 | 3.1 | 16.61 | 23.5 | 31.7 | 50 | 26 | -0.2 | 0.1 | -2 | -1 | PASS | PASS
2023-08-19 15:00 | -0.1 | 16.61 | 23.7 | 31.6 | 51 | 27 | 0.2 | -0.1 | 1 | 1 | PASS | PASS
2023-08-19 16:00 | 0.2 | 16.61 | 23.7 | 31.7 | 50 | 26 | 0 | 0.1 | -1 | -1 | PASS | PASS
2023-08-19 17:00 | -0.6 | 16.6 | 23.9 | 31.8 | 50 | 26 | 0.2 | 0.1 | 0 | 0 | PASS | PASS
2023-08-19 18:00 | 0.9 | 16.61 | 23.9 | 31.8 | 50 | 26 | 0 | 0 | 0 | 0 | PASS | PASS
2023-08-19 19:00 | 0.4 | 16.61 | 24.2 | 32 | 51 | 27 | 0.3 | 0.2 | 1 | 1 | PASS | PASS
2023-08-19 20:00 | -2.1 | 16.62 | 24.3 | 32.1 | 51 | 27 | 0.1 | 0.1 | 0 | 0 | PASS | PASS
2023-08-19 21:00 | -1.9 | 16.61 | 24.2 | 32.1 | 50 | 26 | -0.1 | 0 | -1 | -1 | PASS | PASS
2023-08-19 22:00 | 0.3 | 16.61 | 24.4 | 32.3 | 50 | 26 | 0.2 | 0.2 | 0 | 0 | PASS | PASS
2023-08-19 23:00 | 3.7 | 16.61 | 24.5 | 32.3 | 50 | 26 | 0.1 | 0 | 0 | 0 | PASS | PASS
2023-08-20 | 1.2 | 16.6 | 24.6 | 32.3 | 51 | 27 | 0.1 | 0 | 1 | 1 | PASS | PASS
2023-08-20 01:00 | -0.9 | 16.61 | 24.6 | 32.3 | 52 | 27 | 0 | 0 | 1 | 0 | PASS | PASS
2023-08-20 02:00 | -1 | 16.61 | 24.6 | 32.5 | 52 | 28 | 0 | 0.2 | 0 | 1 | PASS | PASS | Filter Relative Humidity (FRH) < Ambient Relative Humidity (RH) ≈ 25-50% |  |  | PASS
2023-08-20 03:00 | 0.5 | 16.61 | 24.6 | 32.4 | 51 | 27 | 0 | -0.1 | -1 | -1 | PASS | PASS
2023-08-20 04:00 | 0.3 | 16.62 | 24.7 | 32.4 | 51 | 27 | 0.1 | 0 | 0 | 0 | PASS | PASS
2023-08-20 05:00 | 0.7 | 16.6 | 24.8 | 32.5 | 51 | 27 | 0.1 | 0.1 | 0 | 0 | PASS | PASS
2023-08-20 06:00 | 0.6 | 16.6 | 24.5 | 32.4 | 52 | 27 | -0.3 | -0.1 | 1 | 0 | PASS | PASS
2023-08-20 07:00 | 2.2 | 16.62 | 24.7 | 32.3 | 51 | 27 | 0.2 | -0.1 | -1 | 0 | PASS | PASS
2023-08-20 08:00 | 0 | 16.61 | 24.8 | 32.5 | 52 | 27 | 0.1 | 0.2 | 1 | 0 | PASS | PASS
2023-08-20 09:00 | 3.3 | 16.62 | 24.4 | 32.2 | 54 | 28 | -0.4 | -0.3 | 2 | 1 | PASS | PASS
2023-08-20 10:00 | 0.4 | 16.62 | 24.6 | 32.3 | 53 | 28 | 0.2 | 0.1 | -1 | 0 | PASS | PASS
2023-08-20 11:00 | 6 | 16.61 | 24.2 | 32.1 | 54 | 28 | -0.4 | -0.2 | 1 | 0 | PASS | PASS
2023-08-20 12:00 | 0.6 | 16.61 | 24.5 | 32.2 | 53 | 28 | 0.3 | 0.1 | -1 | 0 | PASS | PASS
2023-08-20 13:00 | 1 | 16.61 | 24.4 | 32.1 | 55 | 28 | -0.1 | -0.1 | 2 | 0 | PASS | PASS
2023-08-20 14:00 | 1.7 | 16.61 | 24.4 | 32.1 | 54 | 28 | 0 | 0 | -1 | 0 | PASS | PASS
2023-08-20 15:00 | -0.8 | 16.62 | 24.2 | 31.9 | 58 | 30 | -0.2 | -0.2 | 4 | 2 | PASS | PASS
2023-08-20 16:00 | 4.4 | 16.62 | 24.3 | 32 | 57 | 29 | 0.1 | 0.1 | -1 | -1 | PASS | PASS
2023-08-20 17:00 | 1.3 | 16.62 | 24.2 | 31.8 | 58 | 30 | -0.1 | -0.2 | 1 | 1 | PASS | PASS
2023-08-20 18:00 | 0.7 | 16.61 | 24.7 | 32.2 | 56 | 29 | 0.5 | 0.4 | -2 | -1 | PASS | PASS
2023-08-20 19:00 | 0 | 16.62 | 24.5 | 32.2 | 55 | 29 | -0.2 | 0 | -1 | 0 | PASS | PASS | Flow is continuously stable
2023-08-20 20:00 | -0.4 | 16.62 | 24.7 | 32.3 | 55 | 29 | 0.2 | 0.1 | 0 | 0 | PASS | PASS
2023-08-20 21:00 | 0.6 | 16.61 | 24.7 | 32.4 | 54 | 28 | 0 | 0.1 | -1 | -1 | PASS | PASS
2023-08-20 22:00 | 2.6 | 16.62 | 24.8 | 32.2 | 54 | 29 | 0.1 | -0.2 | 0 | 1 | PASS | PASS
2023-08-20 23:00 | 1.8 | 16.62 | 24.9 | 32.3 | 53 | 28 | 0.1 | 0.1 | -1 | -1 | PASS | PASS
2023-08-21 | 0.3 | 16.63 | 25.2 | 32.3 | 51 | 28 | 0.3 | 0 | -2 | 0 | PASS | PASS
2023-08-21 01:00 | 2.7 | 16.61 | 25.3 | 32.4 | 49 | 26 | 0.1 | 0.1 | -2 | -2 | PASS | PASS
2023-08-21 02:00 | -1 | 16.61 | 24.9 | 32.1 | 48 | 26 | -0.4 | -0.3 | -1 | 0 | PASS | PASS
2023-08-21 03:00 | 0.5 | 16.59 | 25.5 | 32.4 | 46 | 25 | 0.6 | 0.3 | -2 | -1 | PASS | PASS
2023-08-21 04:00 | -0.4 | 16.62 | 25.6 | 32.4 | 45 | 25 | 0.1 | 0 | -1 | 0 | PASS | PASS
2023-08-21 05:00 | 2.2 | 16.61 | 25.6 | 32.4 | 45 | 25 | 0 | 0 | 0 | 0 | PASS | PASS
2023-08-21 06:00 | 0.4 | 16.59 | 25.3 | 32.3 | 45 | 25 | -0.3 | -0.1 | 0 | 0 | PASS | PASS
2023-08-21 07:00 | 2.7 | 16.62 | 24.7 | 31.9 | 48 | 26 | -0.6 | -0.4 | 3 | 1 | PASS | PASS
2023-08-21 08:00 | 1 | 16.62 | 24.8 | 32.1 | 48 | 26 | 0.1 | 0.2 | 0 | 0 | PASS | PASS
2023-08-21 09:00 | 0.2 | 16.61 | 24.2 | 32 | 48 | 25 | -0.6 | -0.1 | 0 | -1 | PASS | PASS
2023-08-21 10:00 | 0 | 16.62 | 24.2 | 32 | 48 | 25 | 0 | 0 | 0 | 0 | PASS | PASS
2023-08-21 11:00 | 0 | 16.61 | 23.7 | 31.6 | 50 | 26 | -0.5 | -0.4 | 2 | 1 | PASS | PASS
2023-08-21 12:00 | 0.2 | 16.61 | 23.6 | 31.4 | 51 | 27 | -0.1 | -0.2 | 1 | 1 | PASS | PASS
2023-08-21 13:00 | 0.5 | 16.61 | 23.7 | 31.4 | 51 | 27 | 0.1 | 0 | 0 | 0 | PASS | PASS
2023-08-21 14:00 | -1.7 | 16.61 | 23.9 | 31.6 | 51 | 27 | 0.2 | 0.2 | 0 | 0 | PASS | PASS
2023-08-21 15:00 | 1.6 | 16.6 | 24 | 31.7 | 51 | 27 | 0.1 | 0.1 | 0 | 0 | PASS | PASS
2023-08-21 16:00 | 2.9 | 16.6 | 24.2 | 31.7 | 50 | 26 | 0.2 | 0 | -1 | -1 | PASS | PASS
2023-08-21 17:00 | 2.7 | 16.6 | 24.2 | 31.8 | 50 | 26 | 0 | 0.1 | 0 | 0 | PASS | PASS
2023-08-21 18:00 | 2.3 | 16.6 | 24.3 | 31.9 | 50 | 26 | 0.1 | 0.1 | 0 | 0 | PASS | PASS
2023-08-21 19:00 | 0 | 16.61 | 24.4 | 32.1 | 49 | 26 | 0.1 | 0.2 | -1 | 0 | PASS | PASS
2023-08-21 20:00 | 0.4 | 16.6 | 24.5 | 32.2 | 50 | 26 | 0.1 | 0.1 | 1 | 0 | PASS | PASS
2023-08-21 21:00 | -1.6 | 16.6 | 24.6 | 32.2 | 50 | 26 | 0.1 | 0 | 0 | 0 | PASS | PASS
2023-08-21 22:00 | 0 | 16.61 | 25 | 32.4 | 50 | 26 | 0.4 | 0.2 | 0 | 0 | PASS | PASS
2023-08-21 23:00 | 0 | 16.62 | 25 | 32.4 | 49 | 26 | 0 | 0 | -1 | 0 | PASS | PASS
2023-08-22 | -1.2 | 16.61 | 25.1 | 32.6 | 48 | 25 | 0.1 | 0.2 | -1 | -1 | PASS | PASS
2023-08-22 01:00 | 3.7 | 16.61 | 25.3 | 32.6 | 48 | 26 | 0.2 | 0 | 0 | 1 | PASS | PASS
2023-08-22 02:00 | 2.5 | 16.6 | 25.3 | 32.7 | 47 | 25 | 0 | 0.1 | -1 | -1 | PASS | PASS
2023-08-22 03:00 | -0.7 | 16.63 | 25.4 | 32.6 | 47 | 25 | 0.1 | -0.1 | 0 | 0 | PASS | PASS
2023-08-22 04:00 | 1 | 16.6 | 25.9 | 32.9 | 45 | 25 | 0.5 | 0.3 | -2 | 0 | PASS | PASS
2023-08-22 05:00 | 2.3 | 16.63 | 25.3 | 32.6 | 46 | 25 | -0.6 | -0.3 | 1 | 0 | PASS | PASS
2023-08-22 06:00 | 1.4 | 16.59 | 26.1 | 32.8 | 45 | 24 | 0.8 | 0.2 | -1 | -1 | PASS | PASS
2023-08-22 07:00 | 0.8 | 16.62 | 25.4 | 32.7 | 46 | 24 | -0.7 | -0.1 | 1 | 0 | PASS | PASS
2023-08-22 08:00 | -0.3 | 16.61 | 24.1 | 32.2 | 47 | 24 | -1.3 | -0.5 | 1 | 0 | PASS | PASS

## Zero Calculation
Time | ConcHR(ug/m3) | Conc(mg) | 1σ | 1σ- | 2σ | 2σ- | Avg Conc | Previous BKGD Value | BKGD Value During Test | Zero Avg (mg) | HR σ | BKGD
2023-08-19 09:00 | 1 | 0.001 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361 | 0 | 0 | 0.00082361 | 0.0015229 | -0.00082361
2023-08-19 10:00 | -0.5 | -0.0005 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 11:00 | 0 | 0 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 12:00 | 1.1 | 0.0011 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 13:00 | 1.6 | 0.0016 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 14:00 | 3.1 | 0.0031 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 15:00 | -0.1 | -0.0001 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 16:00 | 0.2 | 0.0002 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 17:00 | -0.6 | -0.0006 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 18:00 | 0.9 | 0.0009 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 19:00 | 0.4 | 0.0004 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 20:00 | -2.1 | -0.0021 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 21:00 | -1.9 | -0.0019 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 22:00 | 0.3 | 0.0003 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 23:00 | 3.7 | 0.0037 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 | 1.2 | 0.0012 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 01:00 | -0.9 | -0.0009 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 02:00 | -1 | -0.001 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 03:00 | 0.5 | 0.0005 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 04:00 | 0.3 | 0.0003 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 05:00 | 0.7 | 0.0007 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 06:00 | 0.6 | 0.0006 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 07:00 | 2.2 | 0.0022 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 08:00 | 0 | 0 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 09:00 | 3.3 | 0.0033 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 10:00 | 0.4 | 0.0004 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 11:00 | 6 | 0.006 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 12:00 | 0.6 | 0.0006 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 13:00 | 1 | 0.001 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 14:00 | 1.7 | 0.0017 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 15:00 | -0.8 | -0.0008 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 16:00 | 4.4 | 0.0044 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 17:00 | 1.3 | 0.0013 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 18:00 | 0.7 | 0.0007 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 19:00 | 0 | 0 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 20:00 | -0.4 | -0.0004 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 21:00 | 0.6 | 0.0006 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 22:00 | 2.6 | 0.0026 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 23:00 | 1.8 | 0.0018 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 | 0.3 | 0.0003 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 01:00 | 2.7 | 0.0027 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 02:00 | -1 | -0.001 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 03:00 | 0.5 | 0.0005 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 04:00 | -0.4 | -0.0004 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 05:00 | 2.2 | 0.0022 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 06:00 | 0.4 | 0.0004 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 07:00 | 2.7 | 0.0027 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 08:00 | 1 | 0.001 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 09:00 | 0.2 | 0.0002 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 10:00 | 0 | 0 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 11:00 | 0 | 0 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 12:00 | 0.2 | 0.0002 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 13:00 | 0.5 | 0.0005 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 14:00 | -1.7 | -0.0017 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 15:00 | 1.6 | 0.0016 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 16:00 | 2.9 | 0.0029 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 17:00 | 2.7 | 0.0027 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 18:00 | 2.3 | 0.0023 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 19:00 | 0 | 0 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 20:00 | 0.4 | 0.0004 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 21:00 | -1.6 | -0.0016 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 22:00 | 0 | 0 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 23:00 | 0 | 0 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 | -1.2 | -0.0012 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 01:00 | 3.7 | 0.0037 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 02:00 | 2.5 | 0.0025 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 03:00 | -0.7 | -0.0007 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 04:00 | 1 | 0.001 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 05:00 | 2.3 | 0.0023 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 06:00 | 1.4 | 0.0014 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 07:00 | 0.8 | 0.0008 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 08:00 | -0.3 | -0.0003 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361

## Theoretical Zero
Use this sheet to test whether removing the outliers will make an improvement on the BKGD value | [7 empty] | Previous BKGD Value | BKGD Value During Test | Zero Avg (mg) | HR σ | BKGD
Copy values from "Zero Calculations" tab and paste as "Values" | [7 empty] | 0 | 0 | 0.00082361 | 0.0015229 | -0.00082361
Time | ConcHR(ug/m3) | Conc(mg) | 1σ | 1σ- | 2σ | 2σ- | Avg Conc
2023-08-19 09:00 | 1 | 0.001 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 10:00 | -0.5 | -0.0005 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 11:00 | 0 | 0 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 12:00 | 1.1 | 0.0011 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 13:00 | 1.6 | 0.0016 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 14:00 | 3.1 | 0.0031 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 15:00 | -0.1 | -0.0001 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 16:00 | 0.2 | 0.0002 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 17:00 | -0.6 | -0.0006 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 18:00 | 0.9 | 0.0009 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 19:00 | 0.4 | 0.0004 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 20:00 | -2.1 | -0.0021 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 21:00 | -1.9 | -0.0019 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 22:00 | 0.3 | 0.0003 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-19 23:00 | 3.7 | 0.0037 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 | 1.2 | 0.0012 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 01:00 | -0.9 | -0.0009 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 02:00 | -1 | -0.001 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 03:00 | 0.5 | 0.0005 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 04:00 | 0.3 | 0.0003 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 05:00 | 0.7 | 0.0007 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 06:00 | 0.6 | 0.0006 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 07:00 | 2.2 | 0.0022 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 08:00 | 0 | 0 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 09:00 | 3.3 | 0.0033 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 10:00 | 0.4 | 0.0004 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 11:00 | 6 | 0.006 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 12:00 | 0.6 | 0.0006 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 13:00 | 1 | 0.001 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 14:00 | 1.7 | 0.0017 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 15:00 | -0.8 | -0.0008 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 16:00 | 4.4 | 0.0044 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 17:00 | 1.3 | 0.0013 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 18:00 | 0.7 | 0.0007 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 19:00 | 0 | 0 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 20:00 | -0.4 | -0.0004 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 21:00 | 0.6 | 0.0006 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 22:00 | 2.6 | 0.0026 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-20 23:00 | 1.8 | 0.0018 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 | 0.3 | 0.0003 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 01:00 | 2.7 | 0.0027 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 02:00 | -1 | -0.001 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 03:00 | 0.5 | 0.0005 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 04:00 | -0.4 | -0.0004 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 05:00 | 2.2 | 0.0022 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 06:00 | 0.4 | 0.0004 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 07:00 | 2.7 | 0.0027 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 08:00 | 1 | 0.001 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 09:00 | 0.2 | 0.0002 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 10:00 | 0 | 0 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 11:00 | 0 | 0 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 12:00 | 0.2 | 0.0002 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 13:00 | 0.5 | 0.0005 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 14:00 | -1.7 | -0.0017 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 15:00 | 1.6 | 0.0016 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 16:00 | 2.9 | 0.0029 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 17:00 | 2.7 | 0.0027 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 18:00 | 2.3 | 0.0023 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 19:00 | 0 | 0 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 20:00 | 0.4 | 0.0004 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 21:00 | -1.6 | -0.0016 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 22:00 | 0 | 0 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-21 23:00 | 0 | 0 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 | -1.2 | -0.0012 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 01:00 | 3.7 | 0.0037 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 02:00 | 2.5 | 0.0025 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 03:00 | -0.7 | -0.0007 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 04:00 | 1 | 0.001 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 05:00 | 2.3 | 0.0023 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 06:00 | 1.4 | 0.0014 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 07:00 | 0.8 | 0.0008 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361
2023-08-22 08:00 | -0.3 | -0.0003 | 0.0023465 | -0.00069931 | 0.0038695 | -0.0022222 | 0.00082361

## Dropdown (hidden)
Option | Operator | Flow Meter
Yes | Lina Luu | BGI DeltaCal/BX-307 Serial #: 405 Certify Date: 9/5/2023
No | Pam Atcitty | Alicat FP-25 SLPM Serial #: 188862 Certify Date: 8/3/2023
Location
Boneyard
Satellite Office