LEVEL I CHECK PROCEDURES

and

PM2.5 FLOW CHECK PROCEDURE

for

ROLAND AIR MONITORING STATION

Revision: March 2018

Portable equipment to take to site for level I check:

- Laptop computer and site specific memory stick for Roland (memory stick contains level I check forms)

- Perma Pure zero air generator (for ozone level I check) in carrying case

- Thermo 49i-PS O3 calibrator, Thermo 49CPS O3 calibrator, or Environics 6103 ozone (O3) transfer standard

- Sample tubing, zero air tubing with pressure gauge, vent tubing, exhaust tubing, and power cords

Dedicated sampling and calibration instruments used on site:

- ESC 8816 Data System Controller (datalogger)

- Teledyne API 400E Ozone (O3) Analyzer

- Monitor Labs  ML 9841B Nitrogen Oxide (NOx) Analyzer

- Ecotech GasCal 1100 Calibrator

- Ecotech Zero Air Generator, 8301 LC

- Met One PM2.5 BAM 1020

- Cylinder of NO gas (for use in NOx level I check procedure)

Other dedicated equipment on site:

- Bound record book (logbook)

- Plastic wrenches for use in changing sample inlet filters of analyzers

- Teflon sample inlet filter disks (Savillex #1141) for analyzers

- Tweezers for sample inlet filters

- Rolls of filter tape for the BAM

Level I Checks and PM2.5 Flow Check Procedures

for

Roland Air Monitoring Station

The level I check procedure for the ozone analyzer does not use the cylinder of NO gas. Therefore, the level I check of the ozone analyzer can be performed simultaneously with the level I check of NOx analyzer.

- Record your initials, date, and arrival time on site in logbook.

- Turn on laptop computer, plug in memory stick, and open the Roland level I check (aka level I audit) form for the appropriate date. Enter your name, date, and location in the level I check forms.

- If ozone (O3) is offline, as it will be during the winter season, record “offline” in the appropriate blank in level I check form. (Ozone season: March 1 to November 30)

Level I Check Procedure for Ozone (O3)

The level I check procedure for the ozone analyzer is performed independently of the on-site gas calibration system. It utilizes an ozone transfer standard with a zero air generator.

- Fill out the instrument parameters in the level I check form. Use toggle keys for <TST, TST> on API 400E ozone analyzer to scroll through the instrument parameters and obtain values for stability, O3 measure, O3 reference, O3 generator, O3 drive, pressure, sample flow, temperatures (sample, photo lamp, O3 generator, box), slope (gain), and offset.

- Set up zero air generator and ozone generator. Plug power cord of Perma Pure zero air generator into wall outlet and turn zero air generator on. Open rotameter valve all the way. Allow zero air generator ~10 minutes to purge scrubber. If using the Thermo 49i-PS O3 Calibrator transfer standard, use steps A through E. If using the Thermo 49CPS O3 Calibrator transfer standard, use steps F through J. If using the Environics 6103 Ozone Transfer Standard, use steps K through O.

- Plug power cord of Thermo 49i-PS O3 calibrator into the back of the unit and wall outlet.

- Connect one end of the zero air tube with pressure gauge to the zero air generator.

- Remove all of the caps from the back of the Thermo 49i-PS. Connect the vent tube to the “VENT” port and pass the free (open) end of the tube through the exhaust outlet located at the base of the back shelter wall. Also, connect the exhaust tube to the “EXHAUST” port and pass the free end of tube through the exhaust outlet.

- Close the rotameter valve almost completely off on the zero air generator. Power on the Thermo 49i-PS and then connect the zero air tube from the generator to the “IN” port on the back of the transfer standard. Adjust the rotameter valve slowly until the pressure gauge is reading between 10 to 15 psi. Check the pressure occasionally to make sure it doesn’t drift over time.

- Connect one end of the sample tube to the “OZONE” port on the Thermo 49i-PS. Wait for the ozone transfer standard to warm up. The main screen will show a bell emblem on the SET TO line indicating there is an alarm. It may take up to 10-15 minutes for all of the alarms to go away. Once this happens, place the ozone channel into maintenance using step 4.

- Plug power cord of Thermo 49CPS O3 calibrator into the back of the unit and wall outlet.

- Connect one end of the zero air tube with pressure gauge to the zero air generator.

- Remove all of the caps from the back of the Thermo 49CPS. Connect the vent tube to the “VENT” port and pass the free (open) end of the tube through the exhaust outlet located at the base of the back shelter wall. Also, connect the exhaust tube to the “EXHAUST” port and pass the free end of tube through the exhaust outlet.

- Close the rotameter valve almost completely off on the zero air generator. Power on the Thermo 49CPS. Make sure internal pump for Thermo 49CPS comes on. To check, place finger over “EXHAUST” port and feel if air is coming out. If not, remove cover and spin start pump fan (pump should start working). Reattach cover. Then connect the zero air tube from the generator to the “ZERO AIR” port on the back of the transfer standard. Adjust the rotameter valve slowly until the pressure gauge is reading between 10 to 15 psi. Check the pressure occasionally to make sure it doesn’t drift over time.

- Connect one end of sample tube to the “OZONE” port on the Thermo 49CPS. Wait for the ozone transfer standard to warm up. The main screen will show “ALARM” in the lower right hand corner. It may take up to 15-20 minutes for all of the alarms to go away. Once this happens, place the ozone channel into maintenance using step 4.

- Plug power cord of Environics 6103 Ozone Transfer Standard into the back of the unit and wall outlet. Then power the unit on.

- Connect one end of the zero air tube to the zero air generator.

- Remove caps from the “REFERENCE” and “PORT 1” cross fitting, “OUTPUT MANIFOLD PORTS”, “VENT PORT”, and “EXHAUST” port. Leave caps on “PORT 2” through “PORT 5” and on “SAMPLE IN” port. Obtain the large vent hose with 3 small vent tubes. Connect two of the small vent tubes to the “OUTPUT MANIFOLD PORTS” and the other small vent tube to the “VENT PORT” and pass the large vent hose end through the exhaust outlet located at the base of the back shelter wall. Also, connect the exhaust tube to the “EXHAUST” port and pass the free end of tube through the exhaust outlet.

- Press F1 for “CONC” to enter concentration menu, then scroll down to FLOW and enter 5000 using keypad. Press F1 for “START”. Close the rotameter valve almost completely off on the zero air generator. The “REFERENCE” port and “PORT 1” are joined together using a 4-way fitting (cross fitting). Connect a pressure gauge to one of the open inlets on the cross fitting and connect the other end of the zero air tube to the other open inlet on the cross fitting. Adjust the rotameter valve slowly until the pressure gauge is reading between 10 to 20 psi (~ around 15). Check the pressure occasionally to make sure it doesn’t drift over time.

- Connect one end of the sample tube to the last available outlet on the “OUTPUT MANIFOLD PORTS”. Wait for the Environics 6103 to warm up. It may take up to 15-20 minutes before the ozone transfer standard is ready to be used. Once the Environics 6103 is ready, place the ozone channel into maintenance using step 4.

- Login to ESC 8816 Datalogger, if not already. Escape to “Home Menu”. Go to “Login/Set User Level” and enter the password. Password: 13.

- Put channel in maintenance mode. Using ESC datalogger, navigate through “Configuration Menu” to “Configure (Data) Channels” to “Put Channel In Maint.” to place ozone in maintenance (M) mode. Record time (shown on ESC datalogger) in logbook. Escape to “Home Menu” and navigate through “Real-Time Display Menu” to “Display Readings w/ Flags”.

- Disconnect ambient air supply tube (which originates at the glass manifold) from inlet port on back of API 400E ozone analyzer. Connect the sample tube from the ozone transfer standard to “SAMPLE” port on back of ozone analyzer.

- Begin zero point acquisition of level I check. Record start time (as shown on ESC datalogger) in logbook. Wait for ozone concentration values shown on API 400E to stabilize at or near zero.

- While waiting for ozone concentration to stabilize, record which ozone transfer standard and zero air generator are being used to perform the level I audit in the level I check form.

- Record the stable zero concentration for O3 shown on “Real-Time Engineering Flags” screen on ESC datalogger in zero audit blank on level I check form. Record expected ozone concentration shown on ozone transfer standard on level I check form. [Concentration has stabilized when the stability value shown on API 400E is less than 5 ppb.]

- Begin “span” phase of level I check for ozone.

- Thermo 49i-PS – Press ► till the “SET TO” line shows 82 ppb (LEVEL 2), and then press enter. Record time (shown on ESC datalogger) in logbook.

- Thermo 49CPS – Press RUN till the “SET TO” line shows 82 ppb (LEVEL 1), and then press ENTER. Record time (shown on ESC datalogger) in logbook.

- Environics 6103 – Move cursor to O3 using arrow keys. Press F1 for “EDIT” and enter 0.082 ppm using keypad, and then press F1 for “UPDATE”. Record time (shown on ESC datalogger) in logbook.

- Wait for span concentration value on API 400E to stabilize at or near the expected value shown on level I check form. Record stable span concentration of O3 shown on “Real-Time Engineering Flags” screen on ESC datalogger in span audit blank on level I check form. Record expected ozone concentration shown on ozone transfer standard on level I check form.

- Begin “precision” phase of level I check of ozone.

- Thermo 49i-PS – Press ► till “SET TO” line shows 30 ppb (LEVEL 5), and then press enter. Record time (shown on ESC datalogger) in logbook.

- Thermo 49CPS – Press RUN till the “SET TO” line shows 30 ppb (LEVEL 2), and then press ENTER. Record time (shown on ESC datalogger) in logbook.

- Environics 6103 – Move cursor to O3 using arrow keys. Press F1 for “EDIT” and enter 0.030 ppm using keypad, and then press F1 for “UPDATE”. Record time (shown on ESC datalogger) in logbook.

- Wait for precision concentration value shown on API 400E to stabilize at or near the expected value shown on level I check form. Record stable precision concentration of O3 shown on “Real-Time Engineering Flags” screen on ESC datalogger in precision audit blank on level I check form. Record expected ozone concentration shown on ozone transfer standard on level I check form.

- Purge analyzer with pure air.

- Thermo 49i-PS – Press ► till the “SET TO” line shows 0 ppb (ZERO), and then press enter. Record time (shown on ESC datalogger) in logbook. Allow zero air to pass through API 400E until concentration has stabilized at or near zero to purge the system. Level I check of O3 analyzer has now been completed.

- Thermo 49CPS – Press RUN till the “SET TO” line shows 0 ppb (ZERO), and then press ENTER. Record time (shown on ESC datalogger) in logbook. Allow zero to pass through API 400E until concentration has stabilized at or near zero to purge the system. Level I check of O3 analyzer has now been completed.

- Environics 6103 – Move cursor to O3 using arrow keys. Press F1 for “EDIT” and enter 0.000 ppm using keypad, and then press F1 for “UPDATE”. Record time (shown on ESC datalogger) in logbook. Allow zero to pass through API 400E until concentration has stabilized at or near zero to purge the system. Level I check of O3 analyzer has now been completed.

- Replace sample inlet filter, if necessary. The sample inlet filter of the ozone analyzer should be replaced at after the first level I check of each month. The sample inlet filter on the API 400E ozone analyzer is located inside the front cover of the analyzer. Replace the filter as described in steps A through F below.

- Open front cover of analyzer by pulling outward on black plastic locking pins at top of cover.

- Unscrew the metal ring on the filter housing. Remove the metal ring and the underlying glass plate and rubber o-ring. Remove the white teflon ring from the underlying filter.

- Remove old teflon filter. Using tweezers place new teflon filter (Savillex #1141) into filter housing. Make sure filter isn’t wrinkled or torn.

- Place white teflon on top of new filter. Notches on teflon ring must be showing (facing up, away from the filter) to allow sample air to flow through filter properly.

- Place rubber o-ring back in filter housing, and then lay glass plate on top of o-ring. Screw outer metal ring back onto filter housing (over glass plate) tightly, but do not force.

- Close analyzer cover and push black plastic locking pins back into place.

Record sample inlet filter replacement was performed on level I check form and in logbook.

- When the ozone concentration (shown on the API 400E) has stabilized at or near zero, disconnect the sample tube coming from the ozone transfer standard from the “SAMPLE” port on the back of the ozone analyzer. Connect the ambient air supply tube (originating from the glass manifold) to the “SAMPLE” port on back of ozone analyzer.

- Turn off the ozone transfer standard and the zero air generator once sample tubing has been disconnected from the API 400E.

- Take channel out of maintenance mode. Wait for concentration of O3 shown on API 400E to stabilize at ambient readings. Using ESC datalogger, escape to “Home Menu” and navigate through “Configuration Menu” to “Configure (Data) Channels” to “Take Channel Out of Maint.” to take ozone out of maintenance (M) mode. Record the time in logbook. Use time shown on ESC datalogger. Escape to “Home Menu” and navigate to “Real-Time Display Menu” to “Display Readings w/ Flags”.

- Disconnect tubes and power cords from ozone transfer standard and zero air generator. Put the caps back on the ports of the ozone transfer standard.

Level I Check Procedure for NO, NO2, and NOx and PM2.5

The level I check for NOx analyzer uses the NO gas cylinder. Ecotech GasCal 1100 calibrator shows concentration of NO gas being sent through NOx analyzer during each phase (span, precision, titration) of level I check.

- Fill out the instrument parameters in the level I check form. The instrument gain is shown on the main screen of the ML 9841B NOx analyzer. Press “Select” to enter the MAIN MENU. Navigate from MAIN MENU to each of the following menus to find the instrument parameters. [Note: Use “Select” button to enter the menu screens and use “Pg Up” button to return to the previous screen. Using the “Exit” button will return to the analyzer’s main screen.]

- Navigate to INSTRUMENT MENU to MEASUREMENT MENU to find the offsets. Press “Pg Up” twice to return to MAIN MENU.

- Navigate to TEST MENU to OUTPUT TEST MENU to

PREPROCESSOR POTS to locate the instrument’s potentiometer values (Input, High Voltage Adjust). Press “Pg Up” three times to return to

MAIN MENU.

- Navigate to INSTRUMENT STATUS to find the gas flow, gas pressure, ambient pressure, concentration voltage, analog supply, digital supply, ground offset, high voltage, and firmware version. Press “Pg Up” to return to MAIN MENU.

- Navigate to SYSTEM TEMPERATURES to find the instrument’s temperatures (Cell, Converter, Chassis, Manifold, Cooler). Press “Exit” to return to analyzer’s main screen.

- Login to ESC 8816 Datalogger, if not already. Escape to “Home Menu”. Go to “Login/Set User Level” and enter the password. Password: 13.

- Put channels in maintenance mode. Using ESC datalogger, escape to “Home Menu” and navigate to “Configuration Menu” to “Configure (Data) Channels” to “Put Channel In Maint.” to put NO, NOx, and NO2 in maintenance (M) mode. Record the time in logbook. Use the time shown on the ESC datalogger. Escape to “Home Menu” and navigate to “Real-Time Display Menu” to “Display Readings w/ Flags”.

- Begin zero point acquisition of level I check. Using ESC datalogger, escape from “Real-Time Engineering Flags” screen to “Real-Time Display Menu” and navigate to “Display Digital Outputs” and select ZERO and NOX VALVE. This will begin flow of zero (pure) air from the Ecotech GasCal 1100 calibrator. [ESC datalogger display will show that electrical circuits for ZERO and NOX VALVE are “CLOSED”.] Make sure ML 9841B analyzer shows “span measure” on its main screen. Record time shown on ESC datalogger in logbook.

- Wait for zero concentration values shown on ML 9841B analyzer to stabilize at or near the expected values. [Concentration has stabilized when values for instantaneous concentration and time-weighted concentration (average concentration over time interval), as shown on ML 9841B, are the same.] Using GasCal 1100 calibrator, record the flow value for the zero point on the level I check form. Record stable zero concentration values for NO, NOx, and NO2 shown on “Real-Time Engineering Flags” screen of ESC datalogger in zero audit blanks on level I check form.

- Begin span phase of level I check. Using ESC datalogger, navigate from “Real-Time Engineering Flags” screen to “Real-Time Display Menu” to “Display Digital Outputs” and select NOX SPAN to initiate flow of NO gas through the ML 9841B analyzer. Then move cursor to ZERO and press ENTER to stop zero air phase of level I check. [ESC datalogger display will now show that NOX SPAN electrical circuit is “CLOSED” while the ZERO circuit is “OPEN”.] Record time shown on ESC datalogger in logbook. Escape to “Real-Time Display Menu” and navigate to “Display Readings w/ Flags”.

- Verify the cylinder number, concentration, and expiration date of the gas cylinder. Also, check psi values on regulator (A) and cylinder (B) gauges of regulator on NO gas cylinder and record these psi values in level I check form.

- Wait for span concentration values shown on ML 9841B analyzer to stabilize at or near the expected values. Using GasCal, scroll down on main screen to “ZERO FLOW” and “SOURCE FLOW” values. Record flow values for the span point on level I check form. Record stable span concentration values for NO, NOx, and NO2 shown on “Real-Time Engineering Flags” screen of ESC datalogger in span audit blanks on level I check form.

- Begin precision phase on level I check. Using ESC datalogger, navigate from “Real-Time Engineering Flags” screen to “Real-Time Display Menu” to “Display Digital Outputs” and select NOX PRECISION to change concentration of NO gas flowing through ML 9841B analyzer. Then move cursor to NOX SPAN and press ENTER. [ESC datalogger display will now show that NOX PRECISION electrical circuit is “CLOSED” while the NOX SPAN circuit is “OPEN”.] Record time shown on ESC datalogger in logbook. Escape to “Real-Time Display Menu” and navigate to “Display Readings w/ Flags”.

- Wait for precision concentration values shown on ML 9841B analyzer to stabilize at or near the expected values. Using GasCal, scroll down on main screen to “ZERO FLOW” and “SOURCE FLOW” values. Record flow values for the precision point on level I check form. Record stable precision concentration values for NO, NOx, and NO2 shown on “Real-Time Engineering Flags” screen of ESC datalogger in precision audit blanks on level I check form.

- Begin titration (challenge) phase of level I check. Using ESC datalogger, navigate from “Real-Time Engineering Flags” screen to “Real-Time Display Menu” to “Display Digital Outputs” and select NO2 TITRATION to change concentration of NO gas flowing through ML 9841B analyzer. Then move cursor to NOX PRECISION and press ENTER. [ESC datalogger display will now show that NO2 TITRATION electrical circuit is “CLOSED” while the NOX PRECISION circuit is “OPEN”.] Record time shown on ESC datalogger in logbook. Escape to “Real-Time Display Menu” and navigate to “Display Readings w/ Flags”.

- Wait for titration (challenge) concentration values shown on ML 9841B analyzer to stabilize at or near the expected values. Using GasCal, observe the “FLOW” value then scroll down to “SOURCE FLOW” value and record flow values for the titration (challenge) point on level I check form. Record stable challenge concentration values for NO, NOx, and NO2 shown on “Real-Time Engineering Flags” screen of ESC datalogger in challenge blanks on level I check form.

- Using tag on NO gas cylinder, subtract NO concentration shown in ppm from total oxides of nitrogen shown in ppm to get “NO2 impurity” concentration in ppm. Record the NO2 impurity value in level I check form.

- Perform second half of titration phase of NOx level I check. Use Ecotech GasCal 1100 calibrator to send 0.080 ppm of ozone (check concentration of O3 on level I check form) through ML 9841B analyzer to complete titration phase of level I check as follows:

  - Press “Select” to enter main menu. Scroll down to SETUP and press “Select”. Press “Select” again to go into POINTS menu.

  - Scroll down to POINT 4 and press “Select”. Scroll down to P4 O3 CONC and press “Select”.

  - Change ozone concentration from 0 ppm to 0.080 ppm. Use “Select” to move to next position and use “Pg Up” to backspace. Press “Accept”. Press “Exit” to return to main screen. Record time shown on ESC datalogger in logbook.

  - At main screen, scroll down to O3, and make sure that screen shows proper concentration of ozone following through GasCal 1100 to ML 9841B analyzer.

- Wait for titration concentration values shown on ML 9841B analyzer to stabilize at or near expected values. Record stable titration concentration values for NO, NOx, and NO2 shown on “Real-Time Engineering Flags” screen of ESC datalogger in titration blanks on level I check form.

- Stop titration phase of NOx level I check.  Using ESC datalogger, navigate from “Real-Time Engineering Flags” screen to “Real-Time Display Menu” to “Display Digital Outputs” and select ZERO. Then move cursor to NO2 TITRATION and press ENTER to stop flow of NO gas through the ML 9841B analyzer. [ESC datalogger display will now show that ZERO electrical circuit is “CLOSED” while the NO2 TITRATION circuit is “OPEN”.] Record time shown on ESC datalogger in logbook. Escape to “Real-Time Display Menu” and navigate to “Display Readings w/ Flags”. Level I check of NOx analyzer has now been completed.

- Return challenge point to original setting. Using GasCal, repeat steps 13A through 13C to change ozone concentration from 0.080 ppm back to 0 ppm.

- Replace sample inlet filters, if necessary. The sample inlet filter of ML 9841B analyzer should be replaced after the first level I check of each month. The sample inlet filter may be located either at the back of the analyzer or inside the front cover of the analyzer. If sample inlet filter is at back of analyzer, replace it as described in steps A through C below. If sample inlet filter is front cover of analyzer, replace it as described in steps D through G below.

  - Using green plastic wrenches (which fit ribs on outside of inlet filter housing), place one wrench on orange or black half of housing and another wrench on clear half of housing. Twist wrench on orange or black half of housing counterclockwise to loosen housing. Unscrew the two halves of the housing and take them apart.

  - Remove old teflon filter from housing. Using tweezers, place new teflon filter (Savillex #1141) onto plastic screen in orange or black half of inlet filter housing. Carefully seat inner part of clear half of filter housing onto filter, making sure filter does not wrinkle or tear.

  - Screw two halves of filter housing back together. Tighten with green plastic wrenches until snug, but do not force.

  - If sample inlet filter is inside front cover of analyzer, open front cover of analyzer by pulling outward and downward on it from the top.

  - Unscrew metal cap on front of inlet filter housing and pull cap and plunger away from filter.

  - Remove rubber o-ring and old teflon filter from filter housing. Using tweezers, insert new teflon filter (Savillex #1141) into filter housing. Make sure filter isn’t wrinkled or torn. Insert rubber o-ring on top of new filter in housing.

  - Securely seat plunger onto o-ring. Screw on metal cap until hand tight. Close front cover of analyzer.

Record sample inlet filter replacement was performed on level I check form and in logbook.

      - Restore flow of ambient air to analyzer. Wait for concentrations of NO, NOx, and NO2 shown on ML 9841B analyzer to stabilize at or near zero. Then, using ESC datalogger, navigate from “Real-Time Engineering Flags” screen to “Real-Time Display Menu” to “Display Digital Outputs” screen and select NOX VALVE and press ENTER to stop flow of zero air and resume flow of ambient air through the analyzer. Also, select ZERO and press ENTER to stop zero air generator.

      - Wait for concentrations of NO, NOx, and NO2 shown on ML 9841B analyzer to stabilize at ambient air values.

      - Take channels out of maintenance mode. Using ESC datalogger, escape to “Home Menu” and navigate to “Configuration Menu” to “Configure (Data) Channels” to “Take Channel Out of Maint.” to take NO, NOx, and NO2 out of maintenance (M) mode. Record time shown on the ESC datalogger in logbook. Escape to “Home Menu” and navigate to “Real-Time Display Menu” to “Display Readings w/ Flags”.

      - Flow check of PM2.5 analyzer. Using the PM2.5 BAM 1020, navigate from main screen to OPERATE MODE to NORMAL MODE and observe the flow rate while the instrument is sampling. Record flow value in the displayed flow audit blank on PM2.5 BAM 1020 form. The expected flow value is 16.7 LPM.

      - Replace microbalance filter in PM2.5 BAM sensor, if necessary. Open front compartment and inspect the amount of filter tape present on the supply spool (right side). If the supply spool has less than 1/8” filter tape, then replace the filter tape. Refer to Section 3.4 Filter Tape Loading in the BAM 1020 Operation Manual for instructions. After filter tape replacement, perform Self-Test (Section 3.5) and leak check (Section 5.3). Record PM2.5 BAM filter tape replacement was performed on level I check form and in logbook.

Completion of Level I Check and PM2.5 Flow Check Procedures

for

Roland Air Monitoring Station

After completion of level I check (O3, NO, NOx, NO2) and PM2.5 flow check procedures perform the following steps.

- Save level I check forms to Roland memory stick.

- Power down and pack up the computer.

- Pack up the ozone transfer standard and zero air generator including power cords and all of the tubing (zero, sample, vent, exhaust). Check the back of the ozone transfer standard to make sure all of the ports have been capped.

- Check to make sure everything put into maintenance mode has been taken out of maintenance and placed back online. Check instruments and make sure they are all functioning properly.

- Fill out the site checklist. Vacuum site. Record departure time and sign the logbook.