## BAM 1020
| Cherokee Nation Environmental Programs | [11 empty] | BAM 1020 Calibration Report
Site |  |  |  | Stilwell Ncore | [5 empty] | Calibration Performed by |  |  |  | xxxx
Instrument |  |  |  | MetOne BAM 1020 | [5 empty] | Date |  |  |  | 2015-03-12
ID No. |  |  |  | K1143 | [5 empty] | Location |  |  |  | on site
Calibration Equipment
[5 empty] | Model |  |  |  | ID
Flow Calibrator |  |  |  |  | BGI DeltaCal |  |  |  | #1 - SN: 88
Temperature Sensor |  |  |  |  | BGI DeltaCal |  |  |  | #1 - SN: 88
Pressure Sensor |  |  |  |  | BGI DeltaCal |  |  |  | #1 - SN: 88
RH Sensor |  |  |  |  | n/a |  |  |  | n/a
Initial Audit
Size Selective Inlet |  |  |  |  | PM10 |  |  | 2
Flow type |  |  |  |  | Volumetric |  |  | 2
External Ambient sensor |  |  |  |  | BX-596
|  |  | Displayed |  |  | Measured |  |  | Error |  |  | PASS / FAIL
Flow Audit |  |  | 16.7 |  |  | 16.76 |  |  | 0.3593% |  |  | PASS |  | Flow must be within 5%
Ambient temp |  |  | 18.1 |  |  | 18.5 |  |  | -0.4 |  |  | PASS |  | Must be within + 2.5 °C
Ambient BP |  |  | 738 |  |  | 738 |  |  | 0 |  |  | PASS |  | Must be within + 10
|  |  |  | Displayed flow |  |  | Pass / Fail
Leak Check |  |  |  | 0.5 |  |  | PASS
* Clean Nozzle and Vane
Final Leak Check |  |  |  | n/a |  |  | #VALUE!
Self Test |  |  |  | PASS |  |  | 1
Flow Calibration
[5 empty] | Displayed |  |  | Measured
Ambient Temperature | [10 empty] | PASS
Ambient Pressure (Torr) | [10 empty] | PASS
Flow 1 |  | 15
Flow 2 |  | 18.4
Flow 3 |  | 16.7
# Exit calibration menu before performing post calibration check.
Calibration (6 Monthly)
Filter Sensor Calibration
|  |  |  | Displayed |  |  | Measured |  |  | Error
Filter Temperature | [12 empty] | PASS
Filter RH | [9 empty] | 0 |  |  | PASS
Membrane Test
Factory Calibration | [7 empty] | PASS / FAIL |  | #DIV/0!
Test Result
% Difference |  |  |  | #DIV/0!
Post Calibration check
[5 empty] | Displayed |  |  | Measured
Ambient Temperature | [10 empty] | PASS
Ambient Pressure (Torr) | [10 empty] | PASS
Flow 1 |  | 15 | [8 empty] | PASS
Flow 2 |  | 18.4 | [8 empty] | PASS
Flow 3 |  | 16.7 | [8 empty] | PASS
* Flows must be within 0.2 lpm
Notes:
Technicians Signature | [10 empty] | Date: