## NOy
EC9843 NOy Analyzer Calibration Report
Cherokee Nation
Environmental Programs
Site | Stilwell NCore | [7 empty] | Calibration Performed by |  |  |  | J. Adam, D. Keese
Instrument | EC9841-T | [7 empty] | Date |  |  |  | 2019-05-28
ID No. | 07-0850 | [7 empty] | Location (on site or remote) |  |  |  | on site
Calibration Equipment
Calibrator Model |  | Serinus Cal 2000 | [6 empty] | Gas Std. Cylinder Serial No. |  |  |  |  | EB0002834
ID/Serial No. |  | 17-0249 | [6 empty] | Cylinder Concentration (ppm) |  |  |  |  | 27.42
Zero Source Model |  | 8301-LC | [6 empty] | Cylinder Expiry Date |  |  |  |  | 2022-04-13
ID/Serial No. |  | 10-0067* | [6 empty] | Cylinder Pressure (PSI) |  |  |  |  | 1475
Displayed Instrument Parameters
Preprocessor Input |  |  | 40 |  |  | 40 |  |  | High Voltage (V) |  |  |  | 600-700 |  | 643
Preprocessor HV Adjust |  |  | 50-60 |  |  | 54 |  |  | Firmware Version |  |  |  | 1.56.0000 Trace
Gas Flow (slpm) |  |  | 0.45-0.70 |  |  | 0.61 |  |  | Cell Temperature (°C) |  |  |  | 45-55 |  | 49
Gas Pressure (torr) |  |  | 75.0-300.0 |  |  | 111.3 |  |  | Converter Temp. (°C) |  |  |  | 360-370 |  | 361-383**
Ambient pressure 1* (torr) |  |  | 460.0-800.0 |  |  | 715.1 |  |  | Chassis Temperature (°C) |  |  |  | 10-55 |  | 33.9
Ambient pressure 2* (torr) |  |  | 460.0-800.0 |  |  | 723.9 |  |  | Manifold Temperature (°C) |  |  |  | 50-60 |  | 54.8
Analogue Supply (V) |  |  | 11.600-12.200 |  |  | 12.1 |  |  | Cooler Temperature (°C) |  |  |  | 8.0-14.0 |  | 12.3
Digital Supply (V) |  |  | 4.800-5.200 |  |  | 5.1
NOy Converter Efficiency |  |  | 80% - 100% |  |  | 100% |  |  | Instrument Units |  |  | 5 | ppb
Level I Check
|  |  |  | NO |  |  | NOy
Span Instrument Gain |  |  |  | 1.7668 |  |  | 1.7668 |  |  |  | Operating Range (ppb) |  |  |  | 100
Offset (ppb) |  |  |  | 0.128 |  |  | -1.715
[11 empty] | Filter changed? |  | No
Calibration Point (ppb) | Zero flow (slpm) |  | Span flow (slpm) |  | Calculated NOy (ppb) |  |  | Expected (ppb) |  | Measured (ppb) |  |  |  |  | NOy Error (%) | NO Error (%)
[8 empty] | NOy |  | NO |  | NOy
Zero | 7.79 |  | 0 |  | 0 |  |  | 0 |  | 0.2 |  | -1
100 | 4.74 |  | 0.0181 |  | 104 |  |  | 100 |  | 98.4 |  | 99.3 |  |  | -4.519% | -5.385% | < ± 10.1%
20 | 5.75 |  | 0.0043 |  | 20 |  |  | 20 |  | 19.7 |  | 19.2 |  |  | -4% | -1.5% | < ±15.1% or < ± 1.5 ppb
Notes:
GPT (ppb) | NO |  | NOY-NO |  | NOY |  |  | Expected NOY-NO |  | % Drift |  | *ZAG from Pryor. **Temp on converter temp control box is correct.
Challenge | 99 |  | 1.6 |  | 100.6 |  |  | 35.9021 |  | 0.028938
Titration (40 ppb O3) | 63.1 |  | 37.3 |  | 100.5
Source flow | 0.0181
Total flow | 4.75
NO2 impurity | 0.55
Technicians Signature | [8 empty] | 2019-05-28
Technicians Signature | [8 empty] | Date

## SO2
EC9850 SO2 Analyzer Calibration Report
Cherokee Nation
Environmental Programs
Site | Stilwell NCore | [6 empty] | Calibration Performed by |  |  | J. Adam, D. Keese
Instrument | EC9850-T | [6 empty] | Date |  |  | 2019-05-28
ID No. | 07-0757 | [6 empty] | Location (on site or remote) |  |  | on site
Calibration Equipment
Calibrator Model |  |  | Serinus Cal 2000 |  |  |  |  | Gas Std. Cylinder Serial No. |  |  |  | EB0010884
ID/Serial No. |  |  | 17-0249 |  |  |  |  | Cylinder Concentration (ppm) |  |  |  | 8.04
Zero Air Source Model |  |  | 8301-LC |  |  |  |  | Cylinder Expiry Date |  |  |  | 2021-03-26
ID/Serial No. |  |  | 10-0067* |  |  |  |  | Cylinder Pressure (PSI) |  |  |  | 1625
Displayed Instrument Parameters
SO2 Offset (ppb) |  |  |  | 0.2839
Measure Coarse Zero |  |  | 25 - 70 |  | 8 |  |  | Reference Voltage (V) |  |  | 1.00 - 4.00 |  | 2.302
Reference Zero |  |  | 50 |  | 50 |  |  | Analog Supply (V) |  |  | 11.6 - 12.2 |  | 12.1
Measure Gain |  |  | 80 |  | 80 |  |  | Digital Supply (V) |  |  | 4.80 - 5.20 |  | 5.1
Reference Gain |  |  | 5 - 95 |  | 58 |  |  | High Voltage (V) |  |  | 690 - 710 |  | 707
High Voltage Adjust |  |  | 54 -62 |  | 58 |  |  | Lamp Current (mA) |  |  | 34.0 - 40.0 |  | 35.13
Lamp Adjust |  |  | 80 - 90 |  | 83 |  |  | Firmware Version |  |  | 1.35.0002 Trace
Gas Flow (slpm) |  |  | 0.40 - 0.70 |  | 0.58 |  |  | Cell Temperature (°C) |  |  | 47 - 53 |  | 50
Gas Pressure (torr) |  |  | 500 - 800 |  | 676.3 |  |  | Chassis Temperature (°C) |  |  | 15 - 55 |  | 33
Instrument Units |  | 5 | ppb |  |  |  |  | Cooler Temperature (°C) |  |  | 9 - 14 |  | 12.4
Level I Check
Span Instrument Gain |  |  |  | 3.9032 |  | Filter changed? |  |  | No | Full Scale (ppb) |  |  | 100
Calibration scale |  |  |  | 100 | ppb
Calibration Point (ppb) | Zero flow (slpm) |  |  | Span flow (slpm) | Calculated (ppb) |  |  | Expected (ppb) |  | Measured SO2 (ppb) |  | Error
(%)
Zero | 7.75 |  |  | 0 | 0 |  |  | 0 |  | -0.38
52 | 5.71 |  |  | 0.0388 | 54 |  |  | 52 |  | 52.19 |  | -3.352% |  | < ± 10.1%
6 | 7.74 |  |  | 0.006 | 6 |  |  | 6 |  | 5.64 |  | -6% |  | < ±10.1% or < ± 1.5 ppb
Notes:
*ZAG from Pryor.
Technicians Signature | [7 empty] | 2019-05-28

## CO
EC9830 CO Analyzer Calibration Report
Cherokee Nation
Environmental Programs
Site | Stilwell NCore | [5 empty] | Calibration Performed by |  |  |  | J. Adam, D. Keese
Instrument | EC9830-T | [5 empty] | Date |  |  |  | 2019-05-28
ID No. | 07-0839 | [5 empty] | Location (on site or remote) |  |  |  | on site
Calibration Equipment
Calibrator Model |  | Serinus Cal 2000 |  |  |  |  | Gas Std. Cylinder Serial No. |  |  |  |  | EB0048484
ID/Serial No. |  | 17-0249 |  |  |  |  | Cylinder Concentration (ppm) |  |  |  |  | 174.5
Zero Air Source Model |  | 8301-LC |  |  |  |  | Cylinder Expiry Date |  |  |  |  | 2022-03-09
ID/Serial No. |  | 10-0067* |  |  |  |  | Cylinder Pressure (PSI) |  |  |  |  | 1700
Displayed Instrument Parameters
Zero Offset (ppb) |  |  | 13.7961
Measure Coarse Zero |  | 50 - 80 |  | 75 |  |  | Firmware Version |  |  |  | 1.56.0000 Trace
Measure Fine Zero |  | 1-99 |  | 34 |  |  | Cell Temperature (°C) |  |  |  | 47 - 53 |  | 50.8
Input |  | 10-90 |  | 46 |  |  | Converter Temp. (°C) |  |  |  | 87 - 93 |  | 91
Gas Flow (slpm) |  | 0.90 - 1.30 |  | 0.97 |  |  | Chassis Temperature (°C) |  |  |  | 15 - 55 |  | 35.2
Gas Pressure (torr) |  | 430 - 800 |  | 683.5 |  |  | Flow Temperature (°C) |  |  |  | 45 - 55 |  | 49.3
Reference Voltage (V) |  | 3.700 - 4.300 |  | 4.026 |  |  | Cooler Temperature (V) |  |  |  | 2.0 - 2.2 |  | 2.2
Analogue Supply (V) |  | 11.600 - 12.200 |  | 12 |  |  | Mirror Temperature (°C) |  |  |  | 47 - 53 |  | 49.9
Digital Supply (V) |  | 4.800 - 5.200 |  | 5
[7 empty] | Instrument Units |  |  | 5 | ppb
Level I check
Span Instrument Gain |  |  | 0.6114 |  | Filter changed? |  |  | No | Full Scale (ppb) |  |  |  | 600
Calibration scale |  |  | 500 | ppb
Calibration Point | Zero flow (slpm) |  | Span flow (slpm) | Calculated (ppb) |  |  | Expected (ppb) |  | Measured CO (ppb) |  |  | Error
(%)
Zero | 7.79 |  | 0 | 0 |  |  | 0 |  | 24
600 | 4.74 |  | 0.017 | 624 |  |  | 600 |  | 573 |  |  | -8.173% |  | < ± 10.1%
500 | 4.75 |  | 0.0143 | 524 |  |  | 500 |  | 518 |  |  | -1.145% |  | < ±10.1%
170 | 4.75 |  | 0.0049 | 180 |  |  | 170 |  | 188 |  |  | 4.444% |  | < ±10.1%
Notes:
*ZAG from Pryor.
Technicians Signature | [6 empty] | 2019-05-28

## PM 2.5 BAM 1020
| Cherokee Nation Environmental Programs | [10 empty] | BAM 1020 Calibration Report
Site |  |  | Stilwell Ncore | [5 empty] | Calibration Performed by |  |  |  | J. Adam
Instrument |  |  | Met One BAM 1020 | [5 empty] | Date |  |  |  | 2019-05-28
ID No. |  |  | H8930 | [5 empty] | Location |  |  |  | on site
Calibration Equipment
|  |  |  | Model |  |  |  | ID
Flow Calibrator
Temperature Sensor
Pressure Sensor
RH Sensor |  |  |  | n/a |  |  |  | n/a
Initial Audit
Size Selective Inlet |  |  |  | PM2.5 |  |  | 1
Flow type |  |  |  | Volumetric |  |  | 1
External Ambient sensor |  |  |  | BX-596
|  | Displayed |  |  | Expected |  |  | Error |  |  | PASS / FAIL
Flow Audit |  | 16.7 |  |  | 16.7 |  |  | 0% |  |  | PASS |  | Flow must be within ± 5.1%
Ambient temp | [7 empty] | 0 |  |  | PASS |  | Must be within + 2.5 °C
Ambient BP | [7 empty] | 0 |  |  | PASS |  | Must be within + 10
|  |  | Displayed flow |  |  | Pass / Fail
Leak Check | [5 empty] | PASS
* Clean Nozzle, Vane, and VSCC
Final Leak Check | [5 empty] | PASS
Self Test | [5 empty] | 3
Notes:
Technicians Signature | [9 empty] | 2019-05-28

## PM 10 BAM 1020
| Cherokee Nation Environmental Programs | [10 empty] | BAM 1020 Calibration Report
Site |  |  | Stilwell Ncore | [5 empty] | Calibration Performed by |  |  |  | J. Adam
Instrument |  |  | Met One BAM 1020 | [5 empty] | Date |  |  |  | 2019-05-28
ID No. |  |  | K1143 | [5 empty] | Location |  |  |  | on site
Calibration Equipment
|  |  |  | Model |  |  |  | ID
Flow Calibrator
Temperature Sensor
Pressure Sensor
RH Sensor |  |  |  | n/a |  |  |  | n/a
Initial Audit
Size Selective Inlet |  |  |  | PM10 |  |  | 2
Flow type |  |  |  | Volumetric |  |  | 1
External Ambient sensor |  |  |  | BX-596
|  | Displayed |  |  | Expected |  |  | Error |  |  | PASS / FAIL
Flow Audit |  | 16.7 |  |  | 16.7 |  |  | 0% |  |  | PASS |  | Flow must be within ± 5.1%
Ambient temp | [7 empty] | 0 |  |  | PASS |  | Must be within + 2.5 °C
Ambient BP | [7 empty] | 0 |  |  | PASS |  | Must be within + 10
|  |  | Displayed flow |  |  | Pass / Fail
Leak Check | [5 empty] | PASS
* Clean Nozzle, Vane, and VSCC
Final Leak Check | [5 empty] | PASS
Self Test | [5 empty] | 3
Notes:
Technicians Signature | [9 empty] | 2019-05-28