## J8460 Zero Sample
Time | Concentration | 1 Sigma | 1 Sigma- | 2 Sigma | 2 Sigma- | Conc Avg
2009-10-01 18:00 | 0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | Test Records |  |  | Rev C Oct 2011
2009-10-01 19:00 | 0.005 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | BAM-1020 Serial Number: | J4860 |  | Met One Instruments
2009-10-01 20:00 | 0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | Test Performed By: | Dennis Hart |  | Dennis Hart
2009-10-01 21:00 | -0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | Test Start Date: | 2009-10-01
2009-10-01 22:00 | -0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | Test End Date: | 2009-10-04
2009-10-01 23:00 | -0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | Previous BKGD Value: | -0.0014
2009-10-02 | 0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | BKGD Value During Test: | 0
2009-10-02 01:00 | 0 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 02:00 | 0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | Dataset Statistics (milligrams)
2009-10-02 03:00 | 0.003 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | Zero Data Average | 0.0015556
2009-10-02 04:00 | 0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | Hourly Standard Deviation (σ) | 0.0017107 | Within Lab Spec < .0024 mg
2009-10-02 05:00 | 0 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | Background Setting (BKGD) | -0.0015556 | Set this value in the BAM-1020
2009-10-02 06:00 | 0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 07:00 | 0.004 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 08:00 | 0.003 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 09:00 | 0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 10:00 | 0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 11:00 | 0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 12:00 | 0.003 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 13:00 | 0.003 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 14:00 | 0.003 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 15:00 | 0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 16:00 | 0 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 17:00 | 0.003 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 18:00 | 0.004 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 19:00 | 0.004 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 20:00 | 0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 21:00 | -0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 22:00 | 0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-02 23:00 | 0.003 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-03 | 0.003 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-03 01:00 | 0.003 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-03 02:00 | 0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-03 03:00 | 0.003 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-03 04:00 | 0.004 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-03 05:00 | 0 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-03 06:00 | -0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-03 07:00 | -0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-03 08:00 | 0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-03 09:00 | 0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-03 10:00 | 0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-03 11:00 | 0.004 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-03 12:00 | 0.003 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-03 13:00 | -0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | Test Instructions
2009-10-03 14:00 | -0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | 1) The BAM-1020 should be completely installed and operational at its field site before starting a zero test.
2009-10-03 15:00 | 0 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | 2) Allow the BAM-1020 to warm up for several hours before starting a zero test, or plan on discarding the first 4 hours.
2009-10-03 16:00 | 0.003 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | 3) Perform a leak check, nozzle and vane cleaning, and flow check on the unit. See manual section 5.
2009-10-03 17:00 | 0.003 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | 4) In the SETUP > CALIBRATE menu, record the existing BKGD value, then set the BKGD to 0.0000 mg.
2009-10-03 18:00 | 0 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | 5) Remove the PM10 inlet and install the BX-302 zero filter. Run the BAM for at least 72 hours with the zero filter.
2009-10-03 19:00 | 0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | 6) Download the data from the BAM-1020 serial port in .csv format. See manual Section 9.4.
2009-10-03 20:00 | 0 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | 7) Copy the 72 hourly concentration data points into this sheet and analyze the results.
2009-10-03 21:00 | -0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | 8) The negative of the zero average is the correct BKGD (background) setting. Enter it into the BAM-1020.
2009-10-03 22:00 | 0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | 9) Remove the BX-302 zero filter and resume normal BAM operation.
2009-10-03 23:00 | 0.003 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556 | 10) Save the results of this test. The test may be performed annually to audit the performance of the BAM.
2009-10-04 | 0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 01:00 | 0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 02:00 | 0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 03:00 | 0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 04:00 | 0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 05:00 | 0.003 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 06:00 | 0.004 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 07:00 | 0 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 08:00 | -0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 09:00 | 0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 10:00 | 0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 11:00 | 0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 12:00 | 0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 13:00 | 0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 14:00 | 0.002 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 15:00 | 0.004 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 16:00 | 0.004 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556
2009-10-04 17:00 | 0.001 | 0.0032662 | -0.00015513 | 0.0049769 | -0.0018658 | 0.0015556

## Template
Time | Concentration | 1 Sigma | 1 Sigma- | 2 Sigma | 2 Sigma- | Conc Avg
2009-10-01 18:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Test Records
2009-10-01 19:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | BAM-1020 Serial Number:
2009-10-01 20:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Test Performed By:
2009-10-01 21:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Test Start Date:
2009-10-01 22:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Test End Date:
2009-10-01 23:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Previous BKGD Value:
2009-10-02 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | BKGD Value During Test: | 0
2009-10-02 01:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 02:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Dataset Statistics (milligrams)
2009-10-02 03:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Zero Data Average | #DIV/0!
2009-10-02 04:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Hourly Standard Deviation (σ) | #DIV/0! | #DIV/0!
2009-10-02 05:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Background Setting (BKGD) | #DIV/0! | Set this value in the BAM-1020
2009-10-02 06:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 07:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 08:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 09:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 10:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 11:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 12:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 13:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 14:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 15:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 16:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 17:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 18:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 19:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 20:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 21:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 22:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-02 23:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-03 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-03 01:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-03 02:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-03 03:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-03 04:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-03 05:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-03 06:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-03 07:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-03 08:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-03 09:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-03 10:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-03 11:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-03 12:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-03 13:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | Test Instructions
2009-10-03 14:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 1) The BAM-1020 should be completely installed and operational at its field site before starting a zero test.
2009-10-03 15:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 2) Allow the BAM-1020 to warm up for several hours before starting a zero test, or plan on discarding the first 4 hours.
2009-10-03 16:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 3) Perform a leak check, nozzle and vane cleaning, and flow check on the unit. See manual section 5.
2009-10-03 17:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 4) In the SETUP > CALIBRATE menu, record the existing BKGD value, then set the BKGD to 0.0000 mg.
2009-10-03 18:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 5) Remove the PM10 inlet and install the BX-302 zero filter. Run the BAM for at least 72 hours with the zero filter.
2009-10-03 19:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 6) Download the data from the BAM-1020 serial port in .csv format. See manual Section 9.4.
2009-10-03 20:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 7) Copy the 72 hourly concentration data points into this sheet and analyze the results.
2009-10-03 21:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 8) The negative of the zero average is the correct BKGD (background) setting. Enter it into the BAM-1020.
2009-10-03 22:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 9) Remove the BX-302 zero filter and resume normal BAM operation.
2009-10-03 23:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 10) Save the results of this test. The test may be performed annually to audit the performance of the BAM.
2009-10-04 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 01:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 02:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 03:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 04:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 05:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 06:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 07:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 08:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 09:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 10:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 11:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 12:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 13:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 14:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 15:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 16:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
2009-10-04 17:00 |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!

## Raw Data
Time | Concentration