| Colville Confederated Tribes Monitor Network |  |  |  |  |
| --- | --- | --- | --- | --- |
| BAM 1020 Verification Sheet |  |  |  |  |
| Site and Analyzer Information |  |  |  |  |
| Site Name | Nespelem |  | Operator | Kris Ray |
| Site Number | 2 |  | Analyzer S/N | deltaCal 0064 |
| Date |  |  | Last Calibration Date |  |
| Flow Verification (During Monthly Visits) |  |  |  |  |
|  | As Found |  | As Left |  |
| Leak Check |  |  |  |  |
|  |  |  |  | Action ≥ 1.0 L/min Fail ≥1.5 L/min |
|  | Audit Actual |  | BAM 1020 Indicated | Difference |
| Temperature C˚ |  |  |  |  |
|  |  |  |  | Action ± 2.5 C˚ Fail ± 4.0 C˚ |
| Pressure mm Hg |  |  |  |  |
|  |  |  |  | Action ± 6 mmHg Fail ± 10 mm Hg |
| Flow Rate | Actual Flow |  | BAM Indicated Flow | % Difference = (Act Flow - BAM Ind Flow) Actual Flow |
| 16.7 L/min as found |  |  |  |  |
| 16.7 L/min as left |  |  |  |  |
|  |  |  |  | Action ≥ 4% (16.06 to 17.39) Fail ≥ 6% (15.76 - 17.76) |
| 15.0 L/min |  |  |  |  |
| 18.3 L/min |  |  |  |  |
| Self-Test | Pass | Fail | Notes: |  |
| Clean Nozzle/Vane | Yes |  |  |  |
| PM10 Inlet Cleaned | Yes |  |  |  |
| VSCC Cleaned | Yes |  |  |  |
| Data Downloaded | Yes |  |  |  |
| Comments |  |  |  |  |