## Sheet1
BAM 1020
| Date Received: | 2015-09-10 |  | CCID # :
| Date of Test: | 2015-10-19 |  | Serial # : | T20966
| Technician: | XXX
|  |  | Inspect for External Shipping Damage: | Yes
| Any found? | None
|  |  | Inspect for Internal Shipping Damage: | Yes
| Any found? | None
|  |  | Order complete? | NA
List accessories:
| List options: | Ordered for a component spare for the Coarse BAM
Single-point Temperature Verification
| External Reference Standard: | Delta-Cal | S/N: | 1355 | Due Date: | 2015-12-17
| Sensor | Indicated | Reference | Difference | Tolerance | Pass
Ambient Temperature Sensor |  | 24.3 | 24 | 0.3 | ± 2° C | yes
Single-point Barometric Pressure Sensor Verification
| External Reference Standard: | Delta-Cal | S/N: | 1355 | Due Date: | 2015-12-17
| Sensor | Indicated | Reference | Difference | Tolerance | Pass
| Ambient Barometric Pressure Sensor | 697 | 696.5 | 0.5 | ± 10 mmHg | yes
Single-point Flow Rate Verification
| External Reference Standard: | Delta-Cal | S/N: | 1355 | Due Date: | 2015-12-17
| Sensor | Indicated | Reference | Difference | Tolerance | Pass
| Sample Flow | 16.7 | 16.9 | -1.183% | ± 5% | Yes
Leak Check ≤ 1 LPM
| 0.2 | Passed | Yes | Failed
Notes:
| Calibrated temp and BP | Met One | STD
| Temp Before | 22.1 | 24
| Temp As Left | 24.3 | 24
| BP Before | 700 | 696.5
| BP As Left | 697 | 696.5
| Multi Point Flow Verification | 15.17 | 15.18
|  | 18.4 | 18.64
|  | 16.7 | 16.9