## 032516
Model 5014i
| Date Received: | after repair |  | CCID # : | 132985
| Date of Test: | 2016-03-25 |  | Serial # : | CM13361015
| Technician: | XXX and XXX
Inspect for External Shipping Damage:
| Any found? | None
Inspect for Internal Shipping Damage:
| Any found? | None
|  |  | Order complete? | n/a
List accessories:
List options:
Single-point Temperature Verification
| External Reference Standard: | Delta-Cal | S/N: | 428 | Due Date: | 2016-11-09
| Sensor | Indicated | Reference | Difference | Tolerance | Pass
Ambient Temperature Sensor |  | 24.4 | 24.4 | 0 | ± 2° C | yes
| Board Temperature Sensor | 25.2 | 24.2 | 1 | ± 2° C | yes
| Flow Temperature Sensor | 23.8 | 24 | -0.2 | ± 2° C | yes
Single-point Relative Humidity Sensor Verification
| External Reference Standard: | HC2-S | S/N: | 61263434 | Due Date: | 2016-07-20
| Sensor | Indicated | Reference | Difference | Tolerance | Pass
| Ambient RH Sensor | 6.4 | 12.1 | -5.7 | ± 2% | yes
Single-point Barometric Pressure Sensor Verification
| External Reference Standard: | Delta-Cal | S/N: | 428 | Due Date: | 2016-11-09
| Sensor | Indicated | Reference | Difference | Tolerance | Pass
| Ambient Barometric Pressure Sensor | 696.4 | 698 | -1.6 | ± 10 mmHg | yes
Single-point Flow Rate Sensor Verification (Volumetric)
| External Reference Standard: | Delta-Cal | S/N: | 428 | Due Date: | 2016-11-09
| Sensor | Indicated | Reference | Difference | Tolerance | Pass
| Vol.Flow Rate Sensor | 16.67 | 16.44 | 1.399% | ± 4 % | yes
leak test conducted and passed
| External Reference Standard: | Delta-Cal | S/N: | 428 | Due Date: | 2016-11-09
Leak Test - LPM
As Found
Flows measured with NO leak test adapter
| 5014i | 16.67 | -0.23
| Flow standard | 16.44 | 16.67
Flows measured WITH leak test adapter
| 5014i | 16.67
| Flow standard | 16.4 | 16.63
| Measured diff | -0.27
| Corrected diff | -0.04 | Leak check Passed
Mass Calibration
|  | Old Coefficient | 7288.3
|  | New Coefficient | 7140.1
|  | % Difference | -2.0334
|  | Tolerance ± 5% | Pass