## Met Summary
Station Met Summary
Station:
| Previous year | Current year
| verification date | verification date
Ambient Temp Sensor
Verified
Due
Comment
Sonic Anemometer
Verified
Due
Comment
Barometric Pressure Sensor
Verified
Due
Comment
Relative Humidity Sensor
Verified
Due
Comment
Precipitation Gauge
Verified
Due
Comment

## Ambient Temp
Ambient Temperature Sensor Verification
Station:
Date:
Operator:
New temp sensor serial number:
New temp sensor certification date:
Model 045B serial number:
Model 045B date due certification:
As Left
Measurement system verificaton with model 045B calibrator
|  | Switch Setting (Ncore is different) | Known °F | Datalogger Indication °F | Tolerance ± 0.5° F Pass/Fail
|  | T1 8 | -22
|  | T2 9 | 14
|  | T3 10 | 86
|  | T4 11 | 122
As Found
Sensor S/N: |  |  |  | Ambient temp sensor verification:
- Temperature standard -
- Temp standard serial number -
- Date due cert -
- Datalogger indication °F -
- Standard indication °C -
- Standard indication °F -
- Difference °F -
- Pass/Fail ± 1.8° F -
Comment:
076B radiation shield
Date fan replaced:
Comment:

## Wind
50.5 Sonic Anemometer Verification
Station:
Date:
Operator:
Anemometer SN:
| As Found |  | Anemometer firmware number:
Anemometer SN: |  |  |  | As Left
Firmware number: |  |  |  | Firmware 2.62 or newer
Zero test |  |  | Zero test
Bag test - wind speed |  | mph | Bag test - wind speed |  | mph | Acceptable ± 0.45 mph
Span tests |  |  | Span tests
N-S axis blocked |  | mph | N-S axis blocked |  | mph | Acceptable 106.2 to 117.4 mph
N-S axis blocked |  | degrees | N-S axis blocked |  | degrees | Acceptable 7° to 13°
E-W axis blocked |  | mph | E-W axis blocked |  | mph | Acceptable 106.2 to 117.4 mph
E-W axis blocked |  | degrees | E-W axis blocked |  | degrees | Acceptable 157° to 163°
Both axis blocked |  | mph | Both axis blocked |  | mph | Acceptable 106.2 to 117.4 mph
Both axis blocked |  | degrees | Both axis blocked |  | degrees | Acceptable 167° to 173°
| True North orientation |  |  | True North orientation
degrees
[5 empty] | degrees | Acceptable ± 2° of true north
Comment: | [5 empty] | 346°-350°
Comment:

## Baro
Barometric Pressure Verification
Station:
Date:
Operator:
BP model #:
BP sensor SN:
As Found |  |  | As Left
NIST traceable standard |  |  | NIST traceable standard
- Model -
- SN -
- Certification due date -
- Datalogger indication Mbar -
- Standard indication mmHg -
- Standard indication Mbar -
- Difference Mbar -
- Pass/Fail ± 3 Mbar -
Comment:

## Relative Humidity
Relative Humidity Sensor Verification
Station:
Date:
Operator:
New RH sensor model #:
New RH sensor SN:
As Left
Measurement system verificaton with model 045B calibrator
Model 045B serial number:
Model 045B due certification date:
|  | Point | Switch Setting (different at Ncore) | Known % RH | Datalogger % RH | Tolerance ± 1% RH Pass/Fail
|  | (GND) | 3 | 0
|  | V1 | 4 | 50
|  | V2 | 5 | 95
As Found
Sensor S/N: |  |  |  |  | Relative humidity sensor verification:
- Relative humidity stadard model # -
- Relative humidity standard S/N -
- Rel humidity standard cert due date -
- Datalogger indication % RH -
- Relative humidity standard % RH -
- Difference % RH -
- Pass/Fail ±10% RH -
Comments:
Comments:

## Precip
Tipping Bucket Rain Gauge Verification
Station:
Date:
Operator:
| Rain Gauge Model: |  |  |  |  | Rain Gauge Model:
| Rain Gauge SN: |  |  |  |  | Rain Gauge SN:
|  | As Found |  |  |  |  | As Left
Tips at rate of 1 per ~15 sec | Ideal volume of water used (ml) | Actual volume of water used (ml) | Difference | Pass/Fail ± 8.24 ml | Tips at rate of 1 per ~15 sec | Ideal volume of water used (ml) | Actual volume of water used (ml) | Difference | Pass/Fail ± 8.24 ml
10 | 82.4 |  |  |  | 10 | 82.4
| Tips indicated on logger: |  |  |  |  | Tips indicated on logger:
|  |  | Pass/Fail |  |  |  |  | Pass/Fail
Comment:
Comment: