## Explanation
This excel file includes components that may be useful for a Tribal Agency to use in a Level 1 report. The Level 1 Reviewer may often be the Site Operator. The governing QAPP dictates the review process and responsibilities, but all legally defensible data should be reviewed by at least one qualified person independent of the data gathering operations (for the data they review).
The Level 1 Report may include the components of:
1. Title, with Datetime Interval and Parameter
2. Summary of Reviewed and Invalidated Records (with line # in this report and overall completeness so that this report can be quickly understood)
3. Review Criteria, including Expected Range and QAPP Criteria (including Instrument Health)
4. Summary Hourly Plots
5. Specific Intervals Reviewed and Invalidation Decisions, with each interval including:
a. identification of interval and reason data are being scrutinized (generally all L0 flagged records are addressed)
b. questions/decisions regarding any data quality issues
c. screenshot the user can insert if that may be useful for the next reviewer
d. five minute values if available and useful for understanding the data
e. validation/invalidation decision and null codes
6. Reviewer Notes

## Level 1 Report PM10
requires Level 2 approval: | row last edited by: | Level 1 Quarterly Data Review of |  | Jul - Sep 2021 | PM10
Summary:
| aa 11/29/21 | completeness after Level 1 validation= |  | 98.14%
|  | #Invalidated | AQS null code | description |  |  | link to interval | Line # in this sheet
| bb 4/5/21 | 2 | AM | identical consecutive values of 108.782 at 2-3 am on 12/12/20 |  |  | Interval_1 | 71
| bb 4/5/21 | 1 | AM | 1 h extreme negative value of -78.6 on 12/12/20 at 22:00 |  |  | Interval_2 | 92
| bb 4/5/21 | 4 | AN | 4 stuck values at zero at 16:00 on 12/1/20 during site visit |  |  | Interval_3 | 109
|  | 6 | AF | LOST records on 12/31/20 from 18:00 to 23:00 |  |  | Interval_4 | 129
Specific Interval Review Criteria:
|  | Expected range during: | 2021-07-01 | expected minimum = | -2.76 | microg/m3
|  |  | 2021-09-30 23:00 | expected maximum = | 288.01 | microg/m3
| aa 11/29/21 | Documentation that the Site Operator has verified that the maintenance on the equipment and any equipment health issues have been resolved for this time period (see the QAPP criteria checklist in this file) and that the QAPP Review Criteria have been met for this time period (see also the QC charts tab and QAPP Criteria tab)
| aa 11/29/21 | DOCUMENT TYPE | LOCATION | NOTES |  |  | LEVEL | INITIAL
| aa 11/29/21 | Operator logs | Uploaded to QREST | site operator logs extracted from datalogger |  |  | 1 | aa
| aa 11/29/21 | Logbook scans | Uploaded to QREST | images or scans of instrument logbook |  |  | 1 | aa
| aa 11/29/21 | QC check field forms | Uploaded to QREST | images or scans of instrument logbook |  |  | 1 | aa
| aa 11/29/21 | Level 2 reports/assessments | Uploaded to QREST | comparison with [relevant/downwind air station] |  |  | 2 | aa
Hourly Plots with Selected Parameters:
|  | PM10: | hourly | Plotted vs | WS-Resultant
PM10 values outside the range of the Level 0 limits (listed on the next sheet) are in red, values exceeding the Jump criterion are in purple, and Stuck values (N>3) are in gray.
|  | Site Operator comments: | No obvious correlation between PM10 and WS-R, and no correlation with the extreme values of PM10.
bb 4/5/21
|  | PM10: | hourly | Plotted vs | Stnd Deviation of WD
|  | Site Operator comments: | No obvious correlation between PM10 and StdevOfWD, and no correlation with the extreme values of PM10.
bb 4/5/21
|  | PM10: | hourly | Plotted vs | WD
|  | Site Operator comments: | No obvious correlation between PM10 and WD, and no correlation with the extreme values of PM10.
bb 4/5/21
Review of Specific Data Values:
|  | datetime start | datetime end | description
Interval 1: |  | 2020-12-12 02:00 | 2020-12-12 03:00 | Identical values of 108.782 on 12/12 at 2 and 3 am are suspect.
|  | Question for Site Operator? | Was the system preset to run some operation at 2 am?
| bb 4/5/21 | Answer from Site Operator: | No, the system was not preset to run an operation at 2 and 3 am.
bb 4/5/21
|  | 5-minute Values for this Time Interval -> |  | unavailable
QREST Notes on this parameter and interval:
| bb 4/5/21 | No notations for this interval.
Level 1 decision on records nulled:
|  | datetime start | datetime end | AQS Null Code Assigned | # Records Null Coded
| bb 4/5/21 | 2020-12-12 02:00 | 2020-12-12 03:00 | AM | 2
|  | datetime start | datetime end | description
Interval 2: |  | 2020-12-13 22:00 | 2020-12-13 22:00 | One extreme negative value of (-78.6), which occurred at 11 pm so not during a site visit.
|  | Question for Site Operator? | Do you have any information about this one hour?
|  | Answer from Site Operator: | No. There is no explanation for this anomalous value, and it should be invalidated.
|  | 5-minute Values for this Time Interval -> |  | unavailable
QREST Notes on this parameter and interval:
| bb 4/5/21 | No notations for this interval.
Level 1 decision on records nulled:
|  | datetime start | datetime end | AQS Null Code Assigned | # Records Null Coded
|  | 2020-12-13 22:00 | 2020-12-13 22:00 | AM | 1
bb 4/5/21
|  | datetime start | datetime end | description
Interval 3: |  | 2020-12-01 16:00 | 2020-12-01 19:00 | 4 consecutive records of zero, beginning at 16:00 on 12/1/20
|  | Question for Site Operator? | Is there a record of a site visit during this time? Signal was not lost, but 4 consecutive values of exactly zero is not representative of the analyzer's response when it is functioning normally.
| bb 4/5/21 | Answer from Site Operator: | Yes, 12/1/20 at 3pm there was a site visit and resetting the TEOM (turning it off and on) in hopes to clear the high negative PM10 values.
|  | 5-minute Values for this Time Interval -> |  | unavailable
QREST Notes on this parameter and interval:
| bb 4/5/21 | No notations for this interval.
Level 1 decision on records nulled:
|  | 2020-12-01 16:00 | 2020-12-01 19:00 | AN | 4
bb 4/5/21
|  | datetime start | datetime end | description
Interval 4: |  | 2020-12-31 18:00 | 2020-12-31 21:00 | 4 consecutive LOST hours, beginning at 18:00
|  | Question for Site Operator? | Do you have any information about these 4 LOST hours?
| bb 4/5/21 | Answer from Site Operator: | Yes, I had visited the site and exchanged the inlet and got a weekly download and reading.
|  | 5-minute Values for this Time Interval -> |  | unavailable
Level 1 decision on records nulled:
|  | datetime start | datetime end | AQS Null Code Assigned | # Records Null Coded
|  | 2020-12-31 18:00 | 2020-12-31 21:00 | AF | 4
bb 4/5/21
Reviewer Notes:
The above charts of PM10 vs those parameters measured at this site that are likely to influence PM10 concentrations demonstrate no obvious correlation with PM10 overall, and none of these met parameters correspond in any way with the two sets of extreme min/max Level 0-flagged values. The 4 records invalidated due to stuck zero values were invalidated because they are inconsistent with the pattern of the data, and the site operator has no information about these anomalous values.
bb 4/5/21
Note on A2: This column is to notify the Level 2 Reviewer that these are the invalidated or questionable data that should be okayed by the Level 2 reviewer.
Note on C7: blue background means the values in this field is entered by a person and not automatically populated by QREST
Note on F13: QREST could automatically populate this range from the validated data (excluding AQS-nulled values) in QREST, or the Level 1 reviewer can over ride the query and enter into the form what they think the expected range is. (these values are from this site in previous month which is Nov 2020.)
Note on D33: Reviewer can add comments under each chart.
Note on D64: If the Level 1 reviewer enters a question, then that must be answered by the site operator. Could the answer field be red until something is entered into the Answer field?
Note on C75: (if available, 5-min charts zoomed in to the interval being evaluated can be inserted by the reviewer here, with a discussion of whether they provide useful info for justifying validity)
Note on D87: If the Level 1 reviewer enters a question, then that must be answered by the site operator. Could the answer field be red until something is entered into the Answer field?

## OperationalStability
PM10 TEOM 1405 example (thanks to Emma Ruppell and Bishop Paiute); replace with a table relevant to your system.
This form documents the review of logbooks and field data sheets pertinent to the time period and parameter being reviewed. All such original records must be retrievable (paper or scanned copies) and available to an auditor.
Enter date and initials, and location of supporting documentation, ONLY for the operations that were required during the time period being reviewed. Every operation must be verifiable from a scanned or paper logbook page or other source, and that record must be retrievable.
Schedule for Maintenance/Verifications | Daily | Weekly | Bi-weekly | Monthly | Quarterly | Semiannually | Annually | Original Records Available In (complete filename and path or other location):
Check online real-time displays (analyzer and in QREST) for air quality and met data values, check analyzer for codes, flows, and temperatures. | aa 9/27/21 | aa 9/27/21 | aa 9/27/21 | aa 9/27/21 | X | X | X | QREST: scanned logbook pages stored in Supporting Documentation
Verify any alerts- for main flow, status codes, precip, and PM concs; check logger/instrument status for alarms | aa 9/27/21 | aa 9/27/22 | aa 9/27/23 | aa 9/27/24 | X | X | X | QREST: scanned logbook pages
Validate analog data and verify flags |  |  |  | aa 9/27/22 | X | X | X | QREST: scanned logbook pages
Check all registers on instrument and verify that values are within specification. |  | aa 9/27/22 | aa 9/27/22 | aa 9/27/22 | X | X | X | QREST: scanned logbook pages
Physically check flows (total and main), temperature and barometric pressure with calibration device. |  |  | aa 9/27/22 | aa 9/27/22 | X | X | X | QREST: scanned logbook pages
Replace A/C filter and clean condenser coils. Clean A/C filter |  |  |  | aa 9/27/22 | X | X | X | QREST: scanned logbook pages
Check mounting bolts on TEOM enclosure. |  |  |  | aa 9/27/22 | X | X | X | QREST: scanned logbook pages
Clean pump cabinet |  |  |  | aa 9/27/22 | X | X | X | QREST: scanned logbook pages
Replace TEOM filters |  |  |  | aa 9/27/22 | X | X | X | QREST: scanned logbook pages
Exchange inlet with clean inlet. Clean inlet. |  |  |  | aa 9/27/22 | X | X | X | QREST: scanned logbook pages
Measure auxiliary flow with calibration device. |  |  |  |  | X | X | X | QREST: scanned logbook pages
Complete leak check. |  |  |  |  | X | X | X | QREST: scanned logbook pages
Schedule audit with ___?____ and calibrate / repair / replace instrument as needed based on audit findings. | [6 empty] | X | QREST:scanned audit report
Check UPS battery levels. | [5 empty] | X | X | QREST: scanned logbook pages
Check pump (in line). |  |  | aa 9/27/22 | aa 9/27/22 | X | X | X | QREST: scanned logbook pages
Replace in-line filters* |  |  |  |  | X | X | X | QREST: scanned logbook pages
Flow calibration. | [5 empty] | X | X | QREST: scanned logbook pages
Analog calibration. | [6 empty] | X | QREST: scanned logbook pages
Ambient air temperature calibration. | [6 empty] | X | QREST: scanned logbook pages
Ambient pressure calibration. | [6 empty] | X | QREST: scanned logbook pages
Mass transducer calibration verification. | [6 empty] | X | QREST: scanned logbook pages
Replace pump and rebuild | [6 empty] | X | QREST: scanned logbook pages
Clean air inlet | [6 empty] | X | QREST: scanned logbook pages
*During wildfires TEOM filters are replaced more frequently, at total loading of no more than 2200 ug or 90%, as per operational observations
Twice-monthly Checks (records entered from site visits; note that only twice monthly is required for TEOMs, but many programs still check weekly)
Operator | cb | cb | cb | cb
Date | 2021-08-02 | 2021-08-09 | 2021-08-16 | 2021-08-23
Current time (PST) | 9:00 | 10:00 | 11:00 | 12:00
Operating Mode
Status/Codes
Filter Loading %*
Protection Level
Mass Concentration
01-Hr MC
08-Hr MC
12-Hr MC
24-Hr MC
Tot Mass
Case Temp (°C)
Air Temp (°C)
Cap Temp (°C)
Main Flow (Ipm)
Aux Flow (Ipm)
Vac Pressure (atm)
Noise
Frequency
Notes: * Replace TEOM filters when filter loading reaches 90% or every 4 weeks whichever is sooner.

## Level 0 Validation Criteria
Tribal Agency AQS Codes |  |  |  | AQS Site Tribal ID
Meteorological Codes and Level 0 Criteria:
Par Code | Par Name | Method Code | POC | Duration Code | Collect Freq Code | Collection Unit Code | Alert Min | Alert Max | Alert Amt Change | Alert Stuck Count
61103 | Wind Speed - Resultant | 020 | 1 | 1 |  | 012 | 0 |  |  | 3
61106 | Std Dev Hz Wind Direction | 020 | 1 | 1 |  | 014 | 0 | 360 |  | 3
62201 | Relative Humidity | 011 | 1 | 1 |  | 019 | 0 | 100 | 20 | 3
88101 | PM2.5 - Local Conditions | 236 | 1 | 1 | 1 | 105 | -10 | 500 | 100 | 3
62101 | Outdoor Temperature | 040 | 1 | 1 |  | 015 | -60 | 150 | 180 | 3
44201 | Ozone | 047 | 1 | 1 |  | 008 | 0 | 100 | 20 | 3
64101 | Barometric pressure | 011 | 1 | 1 |  | 016 | 0 | 900 | 20 | 3
61104 | Wind Direction - Resultant | 020 | 1 | 1 |  | 014 |  |  |  | 3
88500 | PM2.5 Total Atmospheric | 761 | 3 | 1 |  | 105 | -10 | 5000 | 30 | 3
63301 | Solar radiation | 011 | 1 | 1 |  | 025 | 0 |  |  | 3
62107 | Indoor Temperature | 040 | 1 | 1 |  | 015 | -60 | 150 | 180 | 3
65102 | Rain/melt precipitation | 014 | 1 | 1 |  | 021 | 0 | 0.03 | 0
81102 | PM10 Total 0-10um STP | 079 | 4 | 1 |  | 001 | -10 | 500 | 100 | 3
61101 | Wind Speed - Scalar | 050 | 1 | 1 |  | 012 | 0 | 80 |  | 3
61105 | Peak Wind Gust | 020 | 1 | 1 |  | 012 | 0 |  |  | 3
62103 | Dew Point | 020 | 1 | 1 |  | 015 | -40 |  | 120 | 3

## QAPP Criteria
brown font are example entries from the Site Operator, which are reviewed by the Level 1 and Level 2 reviewers | Excerpted from Element 14 of QAPP (dated______)
Table 14-1. Operational Criteria (must be met for each hourly record):
Enter sufficient information for an auditor to find the evidence and understand the conclusion regarding whether the criterion has been met: |  |  | Table 14-1: Critical Criteria
| Requirement | Frequency |  | Acceptance Criteria | Evidence and Location | Information/Action
3/23/21 verified that at no time in this interval did the analyzer display error codes for out of flow rate status | Average Flow Rate | every 24 hours of op |  | Average within < + 5.1% of design | site logbook, in locked shelter | EPA recommendation
3/23/21 verified that the QC checks conducted on mmddyy and mmddyy passed, and were not more than 30 d apart but more than 14 d apart, that the scanned QC sheets are in Supporting Documentation in QREST, and the QC results were entered into QREST | One-point Flow Rate Verification | every 30 days each separated by 14 days |  | < + 7.1% of transfer standard | See QREST QC record and the Supporting Documentation uploaded by the Site Operator | 1 and 2) 40 CFR Part 58, App A, Sec. 3.3
3) Method 2.10 Table 3-1
Table 14-2. Operational Criteria
Criteria that are important for maintaining and evaluating the quality of the entire data collection system are included in this Operational Criteria Table. Violation of one or more of these requirements may be cause for invalidating a set of data. The decision should consider other quality control information that is available. The dataset for which one or more of these criteria are not met will be considered suspect unless other quality control information demonstrates that the data are valid. The reason for not meeting these criteria MUST be investigated, fixed, or justified.
Table 14-2: Operational Criteria
| Requirement | Frequency | Acceptance Criteria |  | Evidence and Location | Information/Action
mmddyy ____ | System Leak Check | During precalibration check | Auditory inspection with faceplate blocked |  |  | 1, 2 and 3) Method 2.11 Sec. 2.3.2
mmddyy ____ | FR Multi-point | every 365 days and once a | 3 of 4 cal points within < + 10.1% of design |  |  | 1) 40 CFR Part 50 App J Sec. 8.0
mmddyy ____ | Verification/Calibration | calendar year |  |  |  | 2 and 3) Method 2.10 Sec. 2.2.4
mmddyy ____ | Semi Annual Flow Rate Audit | Twice a calendar year and 5- | < + 10.1% of audit standard |  |  | 1, 2) Part 58, App A, Sec. 3.3.3
|  | 7 months apart |  |  |  | 3) Method 2.10 Sec. 7.1.5
mmddyy ____ | Inlet/downtube Cleaning | every 90 days and 4 times a | cleaned |  |  | 1, 2 and 3) Method 2.10 Sec. 6.1.2
calendar year
mmddyy ____ | FR Multi-point Verification | 1/yr. | <± 2.1% |  |  | Part 50, App.L, Sec 9.2
TABLE 14-3. Systematic Criteria
Criteria which are important for the correct interpretation of the data but do not usually impact the validity of a sample or group of samples are included on the third table, the Systematic Issues Table. For example, the data quality objectives are included in this table. If the data quality objectives are not met, this does not invalidate any of the samples but it may impact the error rate associated with the attainment/non-attainment decision.
Table 14-3: Systematic Criteria
| Requirement | Frequency |  | Acceptance Criteria | Evidence and Location | Information/Action
mmddyy ____ | Siting | Every 365 days and 1/ |  | Meets siting criteria or waiver documented |  | 1) 40 CFR Part 58 App E, Sections 2-5
|  | calendar year |  |  |  | 2) Recommendation
3) 40 CFR Part 58 App E, Sections 2-5
mmddyy ____ | Data Completeness | 24-hour |  | > 75% |  | 1, 2 and 3) 40 CFR Part 50 App. K, Sec. 2.3b & c
quarterly
Verification/Calibration Standards and Recertifications - All standards should have multi-point certifications against NIST Traceable standards
mmddyy ____ | Flow Rate Transfer Std. | every 365 days and once a |  | < + 2.1% of NIST-traceable Std. |  | 1) 40 CFR Part 50, App.J Sec. 7.3
|  | calendar year |  |  |  | 2) Method 2.11 Sec. 1.1.3
3) 40 CFR Part 50, App.J Sec. 7.3
mmddyy ____ | Field Thermometer | every 365 days and once a |  | + 0.1degrees C resolution, + 0.1 degrees C accuracy |  | 1, 2 and 3) Method 2.10 Sec. 1.1.2
calendar year
mmddyy ____ | Field Barometer | every 365 days and once a |  | + 1 mm Hg resolution, + 5 mm Hg accuracy |  | 1, 2 and 3) Method 2.10 Sec. 1.1.2
calendar year
mmddyy ____ | Clock/timer Verification | every 180 days and twice a |  | 15 min/day |  | 1) 40 CFR Part 50, App.J Sec. 7.1.5
|  | calendar year |  |  |  | 2) Recommendation
3) 40 CFR Part 50, App.J Sec. 7.1.5

## Notes log PM10
DATA_DTTM | NOTES | NOTE_DATE | USER
9/22/2021 12:00:00 PM | reviewed all intervals | 11/3/2021 12:46:10 AM | bb
9/22/2021 12:00:00 PM | Checked against maintenance records. Checked for internal consistency. | 10/4/2021 5:30:49 PM | aa
9/22/2021 12:00:00 PM | Checked against maintenance records. Checked for internal consistency. | 10/4/2021 5:23:13 PM | aa

## Data PM10
Org ID | Site ID | Par Code | Par Name | Method Code | POC | DateTime (local) | Date (local) | Time (local) | Value | Unit | QREST Flag | AQS Qual Code | AQS Null Code | Lvl1 Review | Lvl2 Review