## CalibDurationEtcCalcs
DURATION CALCULATIONS: |  | (enter values in green cells to see yellow cell values change) |  | DURATION CALCULATIONS: |  | (enter values in green cells to see yellow cell values change)
| enter the CRM's calibration factor in cph per pCi/L [e.g., SunNuclear 1028 cites 2.5 cph/(pCi/L) in literature] |  |  |  | enter the CRM's calibration factor in cph per pCi/L [e.g., SunNuclear 1028 cites 2.5 cph/(pCi/L) in literature]
| CF= | 2 |  |  | CF= | 2
| enter the concentration in pCi/L |  |  |  | enter the concentration in pCi/L
| pCi/L= | 10 |  |  | pCi/L= | 10
| enter the standard error of counts [e.g., enter 0.10 for the ten percent standard error associated with 100 counts] |  |  |  | enter the standard error of counts [e.g., enter 0.10 for the ten percent standard error associated with 100 counts]
| StndError= | 0.03 |  |  | StndError= | 0.032275 |  | tot counts= | 960
| Hours= | 55.5556 | hours exposure needed to produce this standard error in the CRM's counts |  | Hours= | 48 | hours exposure needed to produce this standard error in the CRM's counts
CALIBRATION CHAMBER UNCERTAINTY FOR AVERAGE CONCENTRATION DURING EXPOSURE: |  |  |  | CALIBRATION CHAMBER UNCERTAINTY FOR AVERAGE CONCENTRATION DURING EXPOSURE:
| maximum chamber uncertainty we specify: |  |  |  | maximum chamber uncertainty we specify:
| calibration chamber standard error= | 0.08 | (routine calibration chamber uncertainties including all sources of error, not just counting stats, and LV chamber's published uncertainties for 48 hours was about 6% uncertainty, noting that counting error is NOT the largest error component for chambers) |  | calibration chamber standard error= | 0.08 | (routine calibration chamber uncertainties including all sources of error, not just counting stats, and LV chamber's published uncertainties for 48 hours was about 6% uncertainty, noting that counting error is NOT the largest error component for chambers)
RESULTING CRM UNCERTAINTY COMING OUT OF CALIBRATION: |  |  |  | RESULTING CRM UNCERTAINTY COMING OUT OF CALIBRATION:
|  | 0.08544 | Our goal is a limit of about 12% total uncertainty of a CRM coming out of calibration, because this results in the ability to distinguish between 3.5 and 4 pCi/L for 48-hour measurements, and be correct approximately 84% (1-sigma, 1-tailed) of the time, which requires an approximately 12% 1-sigma total uncertainty at 4 pCi/L. |  |  | 0.086265 | Our goal is a limit of about 12% total uncertainty of a CRM coming out of calibration, because this results in the ability to distinguish between 3.5 and 4 pCi/L for 48-hour measurements, and be correct approximately 84% (1-sigma, 1-tailed) of the time, which requires an approximately 12% 1-sigma total uncertainty at 4 pCi/L.

## OptimalCountingTimeBkgCalib
MITIGATION DECISION optimal counting times:
unit | ratio of indoor test count rate/bkg count rate | cph/(pCi/L) | cph@zero-air | indoor pCi/L | hours_TEST | hours bkg=hours TEST | R(gross)CAL | opt-hours-bkg: | pp 190-191 of Radiotracer Methodology in Biological Science, Wang and Willis, thanks Dr. Jenkins!
femto | 8.4 | 21 | 10 | 4 | 24 | 24 | 84 | 8.2808 |  |  |  | pCi/L | Rg |  | Rb
femto | 5.6 | 21 | 15 | 4 | 48 | 48 | 84 | 20.2837 |  | femto programmable bkg= |  | 25 | 9.78 | cpm/25 pCi/L | 0.25 | cpm
| 7.2 | 18 | 10 | 4 | 24 | 24 | 72 | 8.9443 |  | 0.7 | pCi/L |  | 0.3912 | cpm/pCi/L | 15 | cph
| 7.2 | 18 | 10 | 4 | 24 | 24 | 72 | 8.9443 |  | 14.7 | cph |  | 23.472 | cph/pCi/L
| 7.2 | 18 | 10 | 4 | 24 | 24 | 72 | 8.9443
| 7.2 | 18 | 10 | 4 | 24 | 24 | 72 | 8.9443
| 7.2 | 18 | 10 | 4 | 24 | 24 | 72 | 8.9443
AG | 44.4 | 111 | 10 | 4 | 24 | 24 | 444 | 3.6018
CALIBRATION LAB optimal counting times:
(all units in hours and pCi/L)
unit | ratio of calib chamber count rate/bkg count rate | cph/(pCi/L) | cph@zero-air, or R(bkg) | chamber pCi/L | hours_CAL | hours@zero-air | R(gross)CAL | opt-hours-bkg:
AG | 277.5 | 111 | 10 | 25 | 48 |  | 2775 | 2.8814
femto | 35.7143 | 21 | 14.7 | 25 | 48 |  | 525 | 8.0319
| 45 | 18 | 10 | 25 | 48 |  | 450 | 7.1554
| 31.9444 | 23 | 18 | 25 | 24 |  | 575 | 4.2463
| 16 | 16 | 25 | 25 | 24 |  | 400 | 6
| 25 | 10 | 10 | 25 | 24 |  | 250 | 4.8
| 10 | 4 | 10 | 25 | 24 |  | 100 | 7.5895
Note on D17: bkg programmed in, based on cell K5

## Resultant_CRM_Uncertainty
Change the values in the beige cells, to see the impact of different efficiencies in column B, and change the concentration in the calibrating chamber in column C, to see the impact on the confidence limits of actual field measurements as shown in the yellow columns
| Efficiency (aka CF or sensitivity) | conc | counts accrued for various exposure durations |  |  | 48-hour CRM's counting uncertainty at this concentration | 1-sigma uncertainty of the STAR's avg concentration during 48 hrs | Resultant uncertainty in each CRM calibrated | confidence intervals on the actual field measurement at this concentration and with the CRM uncertainty in column I
Device: | cph/(pCi/L) | pCi/L | counts@1_hr | counts@24_hr | counts@48_hr | (sqrt of counts)/counts | Based on the estimate in UncertaintyIn2ndaryChamber! Sheet, plus 2% to allow chambers some leeway | combining in quadrature cells in coumns G and J | lower bound of 1-sided CI (48 hr) at 84% confidence | upper bound of 1-sided CI (48 hr) at 84% confidence | (one-sided interval so only concerned with lower limit at minus 1 standard error)
A | 2 | 10 | 20 | 480 | 960 | 3.227% | 0.08 | 8.627% | 8.8873 | 11.1127 | '=C4+(SQRT(H4*H4+G4*G4)+J14)*C4
B | 0.6 | 10 | 6 | 144 | 288 | 5.893% | 0.08 | 9.936% | 9.0064 | 10.9936
Sun Nuclear 1027 | 2.5 | 10 | 25 | 600 | 1200 | 2.887% | 0.08 | 8.505% | 9.1495 | 10.8505
Sun Nuclear 1028 | 3 | 10 | 30 | 720 | 1440 | 2.635% | 0.08 | 8.423% | 9.1577 | 10.8423
Sun Nuclear 1029 | 6 | 10 | 60 | 1440 | 2880 | 1.863% | 0.08 | 8.214% | #VALUE! | #VALUE!
1030 Radon Sentinel | 15 | 10 | 150 | 3600 | 7200 | 1.179% | 0.08 | 8.086% | #VALUE! | #VALUE!
femto | 21 | 10 | 210 | 5040 | 10080 | 0.996% | 0.08 | 8.062% | -79.0957 | 99.0957
AlphaGuard | 111 | 10 | 1110 | 26640 | 53280 | 0.4332% | 0.08 | 8.012% | -79.2494 | 99.2494
=SQRT(H4*H4+G4*G4)
[7 empty] | upper bound on BIAS | 0.025
Change the values in the beige cells, to see the impact of different efficiencies in column B, and change the concentration in the calibrating chamber in column C, to see the impact on the confidence limits of actual field measurements as shown in the yellow columns
| Efficiency (aka CF or sensitivity) | conc | counts accrued for various exposure durations |  |  | 24-hour CRM's counting uncertainty at this concentration | 1-sigma uncertainty of the STAR's avg concentration during 48 hrs | Resultant uncertainty in each CRM calibrated | confidence intervals on the actual field measurement at this concentration and with the CRM uncertainty in column I
Device: | cph/(pCi/L) | pCi/L | counts@1_hr | counts@24_hr | counts@48_hr | (sqrt of counts)/counts | Based on the estimate in UncertaintyIn2ndaryChamber! Sheet, plus 2% to allow chambers some leeway | combining in quadrature cells in coumns G and J | lower bound of 1-sided CI (48 hr) at 84% confidence | upper bound of 1-sided CI (48 hr) at 84% confidence | (one-sided interval so only concerned with lower limit at minus 1 standard error)
A | 2 | 10 | 20 | 480 | 960 | 4.564% | 0.08 | 9.211% | 8.8289 | 11.1711 | '=C4+(SQRT(H4*H4+G4*G4)+J14)*C4
B | 0.6 | 10 | 6 | 144 | 288 | 8.333% | 0.08 | 11.55% | 8.8448 | 11.1552
Sun Nuclear 1027 | 2.5 | 10 | 25 | 600 | 1200 | 4.082% | 0.08 | 8.981% | 9.1019 | 10.8981
Sun Nuclear 1028 | 3 | 10 | 30 | 720 | 1440 | 3.727% | 0.08 | 8.825% | 9.1175 | 10.8825
Sun Nuclear 1029 | 6 | 10 | 60 | 1440 | 2880 | 2.635% | 0.08 | 8.423% | 9.1577 | 10.8423
1030 Radon Sentinel | 15 | 10 | 150 | 3600 | 7200 | 1.667% | 0.08 | 8.172% | 9.1828 | 10.8172
femto | 21 | 10 | 210 | 5040 | 10080 | 1.409% | 0.08 | 8.123% | 9.1877 | 10.8123
AlphaGuard | 111 | 10 | 1110 | 26640 | 53280 | 0.6127% | 0.08 | 8.023% | 9.1977 | 10.8023
=SQRT(H4*H4+G4*G4)
[7 empty] | upper bound on BIAS | 0.025
Note on H3: the derived value for LV is 6%, but we may add~2% due to bkg variability
Note on H19: the derived value for LV is 6%, but we may add~2% due to bkg variability

## UncertaintyIn2ndaryChamber
ESTIMATION OF THE UNCERTAINTY OF CALIBRATING CHAMBER'S AVG CONC FOR 48 HOUR EXPOSURE (green cells show the uncertainty in each Lucas cell filled by NAREL from the bubbler [the calibrating lab ships 3 cells to NAREL to get filled])
Blue cells show estimates of the uncertainty of the calibrating chamber's average concentration for 48 hour calibration exposure durations (please see comments explaining some tricky cells)
Component of Uncertainty Rows 4-11 calculate the Primary Standard (NAREL's NIST SRM Ra226 in solution [bubbler]-generated Rn222 gas concentration; note that these values are all based on LV data and NAREL will be different, but not a lot greater, we hope. | Reference or source of estimate | s, measured standardized, or % uncertainty) | N | STANDARD ERROR OF THE MEAN = s/sqrt(N) which is estimate of s of the mean if N>0 | s2 | sum | s | estimated uncertainty, 1 s | Upper Limit of Systematic Errors
SRM 4964-B 226Ra standard bubbler solution | NIST certificate | 0.0012667 | n/a |  | 1.6044e-06 |  |  |  | 0.0012
Volumetric transfer to Lucas cells | scale calibration certificate | 3.3333e-05 | n/a |  | 1.1111e-09 |  |  |  | 0.0001
| manometer calibration certificate | 5e-05 | n/a |  | 2.5e-09 |  |  |  | n/a****
| repeated flushing of laboratory standards (a) | 6e-05 | n/a |  | 3.6e-09 |  |  |  | 6e-05
| elapsed time resolution, 1 min/14 days | 4.9603e-05 | n/a |  | 2.4605e-09 |  |  |  | n/a****
| calculation of fraction of remaining atoms based on vacuum of 30 mtorr (b) | 3.9474e-05 | n/a |  | 1.5582e-09 |  |  |  | 3.9474e-05
Lucas Cell CF generation counting error (since the Lucas cells belong to the calibration lab that ships them to NAREL to be filled, these CFs will be calculated by the calibration lab prior to shipping them to NAREL, and are basically what will be calibrated by this process, then the "authority" of the NIST SRM is carried through via these cells (and thence traceable to) the actual calibration chamber concentration) | counting error of CF generation process (these values based on LV cells counting ~25 pCi/L) | 0.0093487 |  |  | 8.7399e-05 |  |  |  | (below this row, no systematic errors measured or estimated)
SUBTOTAL: cumulative error to this point in the concentration in each of the NAREL-filled Lucas cells is shown in cell I11 | the square root of the sum of the squares of each component standard uncertainty plus each upper limit of systematic errors in rows above |  | 1 |  |  | 8.9014e-05 | 0.0094347 | 1.083%
LUCAS CELLS SENT FROM NAREL have estimated uncertainty in each Lucas cell filled w bubbler, including counting error and other cumulative error to this point, in cell C12=I11, but because 3 Lucas cells are shipped from NAREL to Phil or Jim or other 2ndary chamber this is diminished by the sqrt of 3 to result in the uncertainty of the "known" from NAREL to be as shown in cell F12 |  | 1.083% | 3 | 0.0062551 | 0.003913%
COUNTING of the 3 Lucas cells sent from NAREL filled with bubbler-generated concentration adds an uncertainty due to the difference between the 3 cells as counted by the calibration chamber |  | 0.035 | 3 | 0.020207 | 0.00040833
NOW, the CHAMBER MAKES 3 GRABS w the same LUCAS CELLS (which are theirs anyway as they shipped them to NAREL to get filled), and the uncertainty in the average of these 3 Lucas cell chamber grabs is used to calibrate the chamber conc. by comparing this average to their counts from the cells as filled by NAREL with the "known" bubbler-generated concentration. (Note that is row contributes uncertainty based only on the agreement between the 3 chamber grabs and not including the cumulative error shown in cell F12 which is the uncertainty in the "known" conc based on the average of 3 cells as filled by NAREL, which gets carried though the process.) |  | 0.035 | 3 | 0.020207 | 0.00040833
Phil's chamber uses 2 large flow through scint cells called SACs (as did LV), and each SAC's counting error of the hour when the grab was made, and used generate each SAC's CF (based on 1 hour counting error; 3% is high estimate) is shown in cell C15 as the standard error |  | 0.03 | 1 |  | 0.0009
SAC counting error to generate the 2nd SAC's CF |  | 0.03 | 1 |  | 0.0009
agreement between the set of the SACs for 1 chamber hour |  | 0.02 | 2 | 0.014142 | 0.0002
SUBTOTAL: estimated 1 sigma up to this point for one hour of chamber data from avg of 2 SACs | [5 empty] | 0.0028558 | 0.05344 | 0.05344
48 hour counting error of one SAC |  | 0.007 |  |  | 4.9e-05
48 hour counting error of the paired SAC |  | 0.007 |  |  | 4.9e-05
agreement between the set of the SACs for 48 chamber hours |  | 0.035 | 2 | 0.024749 | 0.0006125
[6 empty] | 0.0035663 | 0.059718 | 5.972% | for 48 hr chamber exposure
**Report of Traceability from NIST shows agreement of -1.61% (1993)
***I am inferring Paul used a parallel process of generating the bubblers to what he did when generating 2557-3 and 2557-4, which is to compare the volumetric transfer with gamma spectroscopy measurements of activity, then calculating specific activity based on the bubbler volume, so that is not 2 replicates but at least two assessments. These values of one percent for imprecision and 1.5 percent for systematic errors are taken from 2557-3 which shows a standard error for imprecision of 1% and systematic error of 3%
****digital prevents bias in recording value
(a) repeated flushing of bubblers: if you repeat the exercise with a second set of five cells, as quickly as possible after the first set of five, you will get the same result. This value was taken from the Swedes' study, but is so low it doesn’t make a difference.
[do we have data we can use for this, where the cell CF generation exercise was repeated? Or, can we use the volume
represented by 30 millitorr of pressure in 0.36 liter?]
(b) this is linearly related to pressure, but also did a calculation based on avogadro's number of 6.023e23 molecules in 22.4 liters at 1 atm, which yields
2.7e19 molecules in one cc at one atm
Comment from Arthur Scott: The operator (and the customer) should have reasonable expectation that a field measurement of 3.5 pCi is indeed less than 4pCi. This implies a total uncertainty (count statistics plus calibration) of <12.5%. How much of this uncertainty can reasonably allocated to calibration?
If the nominal CRM has a sensitivity of 2 cph/pCi/l, then the sd of a 48 h count at 4 pCi is 5%, which would allow for a calibration bias/uncertainty up to about 8%.
The calibration is carried out at a concentration high enough for the statistical error in CRM counts to be about 2%, which leaves around 6% uncertainty for STAR concentration measurement statistical uncertainties plus transfer system biases. For example, if Montgomery says the Lucas cells are good to 3%, there is a potential bias of 2-3%, depending on number of cells used, plus the statistical error in cell and chamber measurement system count etc.
All in all, I think a total calibration uncertainty in the region of 6% is necessary, and achievable, and at the least could be used as a basis for discussion.
Note on D12: 3 Lucas cells shipped from NAREL to calibrating chambers
Note on F12: The avg of 3 diminishes the uncertainty of this "known" concentration in the Lucas cells, counted by the calibrating chamber and used as their "known" against which to calibrate THEIR chamber.
Note on C13: The standard error of the difference between the counts of the 3 grabs made w Lucas cells of the calibrating chamber concentration; the LV chamber agreement between grabs showed an avg standard error of 3%. Given the large number of counts esp if the calibrating chambers concentrations ~>20 pCi/L, agreement in Phil's or Jim's may be less than 3%.
Note on C14: Since these are the same or similar cells, this assumes that the differences between them will be about the same as the differences between the 3 cells that were filled by NAREL and counted at the calibration chamber as shown in cell C13
Note on C17: Agreement between the 2 SACs in LV chamber was 1.5%, and this assumes that Phil's will be similar
Note on C19: estimates in cells C19 and C20 are based on calculations from LV data, but if concentrations are ~>20 pCi/L results in Phil's chamber should be similar

## QC_scribbling
|  |  | limits are 36% and 67% RPE and RPD | [8 empty] | 36% | 67%
CROSSCHECKS |  |  | RPE/COV=sqrt(2) |  |  | QC CHECKS (DUPLICATES)
CRM1 | CRM2 | RPE | COV |  | based on in control of: | CRM1 | CRM2 | RPD | COV | RPE/COV=sqrt(2) | based on in control of:
2 | 1.2 | 66.67% | 47.14% | 1.414% | 0.24051 | 2 | 1 | 66.67% | 47.14% | 1.414% | 0.24051
3.5 | 2.1 | 66.67% | 47.14% | 1.414% | 0.24051 | 3.5 | 1.7 | 69.23% | 48.95% | 1.414% | 0.24976
3.7 | 2.2 | 68.18% | 48.21% | 1.414% | 0.24598 | 3.7 | 1.8 | 69.09% | 48.85% | 1.414% | 0.24926
6 | 4.4 | 36.36% | 25.71% | 1.414% | 0.13119 | 6 | 4.2 | 35.29% | 24.96% | 1.414% | 0.12733
8 | 5.9 | 35.59% | 25.17% | 1.414% | 0.12841 | 8 | 5.6 | 35.29% | 24.96% | 1.414% | 0.12733
12 | 8.8 | 36.36% | 25.71% | 1.414% | 0.13119 | 12 | 8.3 | 36.45% | 25.78% | 1.414% | 0.13151
10 | 7.3 | 36.99% | 26.15% | 1.414% | 0.13344 | 10 | 20 | -66.67% | 47.14% | -1.414% | 0.24051
5 | 3.7 | 35.14% | 24.84% | 1.414% | 0.12676 | 5 | 7 | -33.33% | 23.57% | -1.414% | 0.12026
14 | 10 | 40% | 28.28% | 1.414% | 0.14431 | 14 | 25 | -56.41% | 39.89% | -1.414% | 0.20351
34 | 25 | 36% | 25.46% | 1.414% | 0.12988 | 15 | 25 | -50% | 35.36% | -1.414% | 0.18038
17 | 27 | -37.04% | 26.19% | -1.414% | 0.13362 | 17 | 27 | -45.45% | 32.14% | -1.414% | 0.16399
4 | 8 | -50% | 35.36% | -1.414% | 0.18038
8 | 4 | 100% | 70.71% | 1.414% | 0.36077

## scribbling
[5 empty] | from online device descriptions, per pCi/L: |  |  | AG: 5 cpm at 100 Bq/m³
| value | ln of data |  |  | cpm | cph |  | 0.05 cpm per Bq/m3 |  |  | pCi/L | Bq m3
| 1.1 | 0.09531 |  | AG | 5 | 300 |  | 1.85 | AG | http://www.nrsb.org/images/file/NRSB%20Devices%20Ltg.pdf | 4 | 148
| 1.1 | 0.09531 |  | femto | 0.3 | 18 |  | so in terms of cph |  |  | 9 | 333
| 1.1 | 0.09531 |  | Sun Nuclear 1029 CRM |  | 6 |  | 111 |  |  | 159 | 5883
| 1.1 | 0.09531 |  | 1030 Radon Sentinel |  | 15 |  |  |  | http://www.ludlums.com/component/virtuemart/equipment-type-3/sample-counting-21/general-purpose-scaler-34-detail?Itemid=0 | 1 | 37
| 1.1 | 0.09531 |  | RADALINK Radon Telemonitor | 0.33 | 19.8 |  |  |  | http://www.ludlums.com/multisites/medphys/images/stories/product_manuals/M2000.pdf | 120 | 4440
| 1.1 | 0.09531 |  | Sun Nuclear 1028 CRM |  | 3 |  |  |  | http://ludlums.com/images/stories/data_sheets/M2000.pdf | 73 | 2701
| 2 | 0.69315 |  | 1030 Radon Sentinel |  | 15 |  | 5 cpm at 100 Bq/m³ (3 pCi/l) |  |  | 159 | 5883
| 2 | 0.69315 |  | Sun Nuclear 1027 |  | 2.5 |  |  |  |  | 0.7 | 25.9
| 2 | 0.69315 | [8 empty] | 10 | 370
| 3 | 1.0986 | [8 empty] | 87 | 3219
| 3 | 1.0986 |  | from Andy George's sheet in cpm/4 pCi/L: | cpm/pCi/L | cph/pCi/L |  |  |  |  | 100 | 3700
| 3 | 1.0986 |  | 1.2 | 0.3 | 18 | femto |  |  |  | 50 | 1850
| 4 | 1.3863 |  | 7.6 | 1.9 | 114 | old AG |  |  |  | 20 | 740
| 4 | 1.3863 |  | 0.17 | 0.0425 | 2.55 | Sun Nuclear 1027 |  |  |  | 10 | 370
| 5 | 1.6094 |  | 2 | 0.5 | 30 | Rad 7 |  |  |  | 12 | 444
| 7 | 1.9459 | [8 empty] | 11 | 407
| 15 | 2.7081 | [8 empty] | 5 | 185
stats | ln stats | [9 empty] | 4 | 148
3.3294 | 0.88136 | 2.4142 | [8 empty] | 3 | 111
3.4353 | 0.76914 |  | 2.1579 | geometric stnd deviation | [6 empty] | 2 | 74
2.4142 | [10 empty] | 2.7 | 99.9
| count= | 17 | [8 empty] | 8.1 | 299.7