DOCUMENT TYPE:		Standard Operating Procedure

TITLE:					Cleaning Pump Muffler on BAM 1020

PREPARED BY:  						                        06/23/2011

Tribal Air Program					                              	(Date)

REVIEWED BY:

Environmental Supervisor                                  						(Date)

APPROVED BY:

Environmental Department Manager							         (Date)

Table of Contents

- Scope of Application……………………………………………………………………………………………….1

- Summary of Method………………………………………………………………………………………….…….1

- Interferences/Comments……………………………………………………………………………………………1

- Supplies………………………………………………………………………………………………….………….2

- Procedures Setups………………………………………………………………………..………………………….2

- Troubleshooting…………………………………………………………………………….……………………….2

# Scope of Application

## Gather ambient air data, which have been linked to smog, ozone depletion, acid rain and global warming. Maintaining the equipment that records this data is a vital aspect of providing consistent, current and valid data.

# Summary of Method

## Remove muffler

## Clean muffler

## Reassemble muffler

# Interferences/Comments

## The pump runs almost continuously and thus is hot; be sure to take proper precaution when removing the muffler.

## The pump may be allowed to run while the muffler is removed.

# Supplies

## Adjustable wrench

## Isopropyl alcohol

## Cotton swabs

## Teflon tape

## Small cup

# Procedure

## Remove the muffler by unscrewing the hex-nut counter clockwise.

## After it is cool to the touch, remove the inside carbon build up with a cotton swab.

## Soak the muffler in isopropyl alcohol for 15 minutes, agitating every couple of minutes.

## Clean the inside again using a cotton swab.

## Rinse with alcohol or water.

## Let the muffler set in the sun for 15 minutes to ensure complete dryness.

## Wrap muffler’s threaded end with Teflon tape to make sure threads are sealed.

## Replace muffler at the same tightness as it was remove.

# Troubleshooting

## Contact the manufacturer at www.gastmfg.com