LEVEL I CHECK PROCEDURES

and

PM 10 and PM2.5 FLOW CHECKS PROCEDURES

for

ROLAND AIR MONITORING STATION

Revision 0: 4 October 2010

Portable equipment to take to site for level I check:

- Laptop computer and site specific memory stick for Roland (memory stick contains level I check forms)

- Perma Pure zero air generator (for ozone level I check) in carrying case

- Thermo 165C ozone (O3) generator in carrying case, Thermo 49C O3 Calibrator, or Environics 6103 ozone (O3) transfer standard

- Sample tubing, zero air tubing with pressure gauge, vent tubing, and power cord

Dedicated sampling and calibration instruments used on site:

- ESC 8816 Data System Controller (datalogger)

- Teledyne API 400E Ozone (O3) Analyzer

- Monitor Labs  ML 9841B Nitrogen Oxide (NOx) Analyzer

- Ecotech GasCal 1000 Calibrator

- Ecotech Zero Air Generator, 8301 LC

- R&P Series 1400a TEOM PM2.5 Monitor

- Thermo Anderson PM10 BAM FH62-C14

- Cylinder of NO gas

Other dedicated equipment on site:

- Printer for generating hard copies of level I check report

- Bound record book (logbook)

- Plastic wrenches for use in changing sample inlet filters of analyzers

- Teflon sample inlet filter disks (Savillex #1141) for analyzers

- Tweezers for sample inlet filters

- TEOM microbalance filter disks (TX40 molded filters)

- Fork for inserting TEOM microbalance filter disks

- Roll of filter tape for the BAM

Level I Check Procedure for NO, NO2, and NOx, PM10 and PM2.5*

Level I check for NOx analyzer use the NO gas cylinder. Ecotech GasCal 1000 calibrator shows concentration of NO gas being sent through NOx analyzer during each phase (span, precision, titration) of level I check. * PM10 and PM2.5 are disabled due to discontinuation; however, they can be reactivated at anytime for regulatory monitoring and/or special projects.

- Record your name, date, and arrival time on site in logbook.

- Turn on laptop computer, plug in memory stick, and open the Tahlequah level I check (aka level I audit) file for the appropriate date. Enter your name and the date in the level I check form.

- If ozone (O3) is offline, as it will be in winter season, record “offline” in appropriate blank in level I check form on laptop. (Ozone season: March 1 to November 30.)

- Preliminary checks. Using Monitor Labs ML 9841B NOx analyzer, complete the following series of four preliminary instrument checks.

  - Check “instrument gain” shown on main screen of the instrument. Record gain value in level I check form on laptop.

  - Check to see if there are any warning “flags” (such as “zero flow”) shown on main screen of the analyzer. If a flag is present, then navigate from main menu screen to “event log” screen and/or “system faults” screen to determine if analyzer requires maintenance, calibration, or other adjustment. Describe flag and any associated event or system fault in logbook.

  - Navigate from main menu screen of the instrument through instrument menu screen to measurement menu screen to check instrument noise. Instrument noise should be 50.0 ppb or less. If instrument noise of any of the analyzers exceeds these values, there may be problems with the affected analyzer to be corrected.

  - Navigate from main menu screen through test menu and output test menu screens to “preprocessor pots” screen to check instrument potential values. The test measure value should be zero and the other instrument potential values should be greater than zero and less than 99. If instrument potentials of any of the analyzers exceed these values, there may be problems with the affected analyzer that need to be corrected.

- Put channels in maintenance mode. Using ESC datalogger, navigate through “configuration menu”, “configure (data) channels”, and “put channel in maintenance” screens to put NO, NOx, and NO2 in maintenance (M) mode. Record the time this was done in logbook. Use the time shown on the ESC datalogger.

- Begin zero point acquisition of level I check. Using ESC datalogger, navigate from “real time display w/ flags” screen to “display digital outputs” screen and select each of the following: ZERO; and NOX VALVE. This will send zero (pure) air through the NO analyzer. Make sure ML 9841B analyzer shows “span measure” on its main screen. Record time shown on ESC datalogger in level I check form on laptop and in logbook.

- Wait for zero concentration values shown on ML 9841B NOx analyzer to stabilize at or near the expected values. Record stable zero concentration values for NO, NOx, and NO2 shown on “real time display w/ flags” screen of ESC datalogger in zero audit blanks on level I check form on laptop.

- Begin span phase of level I check. Using ESC datalogger, navigate from “real time display w/ flags” screen to “display digital outputs” screen and select NOX SPAN to initiate flow of nitrogen oxide gas through NOx analyzers. Then move cursor up to ZERO and press ENTER to stop zero air phase of level I check. Record time shown on ESC datalogger in level I check form on laptop and in logbook.

- Verify the cylinder number, concentration, and expiration date of the gas cylinder. Also, check psi values on regulator (A) and cylinder (B) gauges of regulator on NO gas cylinder and record these psi values in level I check form on laptop.

- Wait for span concentration values shown on ML 9841B NOx analyzer to stabilize at or near the expected values. Record stable span concentration values for NO, NOx, and NO2 shown on “real time display w/ flags” screen of ESC datalogger in span audit blanks on level I check form on laptop. Using GasCal 1000 calibrator, use arrow key to scroll down on main screen to “zero flow” and “source flow” values. Record flow values in “dilution calculator” on level I check form on laptop.

- Begin precision phase on level I check. Using ESC datalogger, navigate from “real time display w/ flags” screen to “display digital outputs” screen and select NOX PRECISION to change concentration of NO gas flowing through NOx analyzer. Then move cursor up to NOX SPAN and press ENTER to stop span phase of level I check. In effect, the dilution of NO gas being sent to the analyzer has been changed from the span dilution (span concentration) to the precision dilution. In other words, less NO gas is now being mixed with the zero air that is sent to the analyzer. Record time shown on ESC datalogger in logbook.

- Wait for precision concentration values shown on ML 9841B NOx analyzer to stabilize at or near the expected values. Record stable precision concentration values for NO, NOx, and NO2 shown on “real time display w/ flags” screen of ESC datalogger in precision audit blanks on level I check form on laptop. Using GasCal 1000 calibrator, use arrow key to scroll down on main screen to “zero flow” and “source flow” values. Record flow values in “dilution calculator” on level I check form on laptop.

- Begin titration (challenge) phase of NOx level I check. Using ESC datalogger, navigate from “real time display w/ flags” screen to “display digital outputs” screen and select NO2 TITRATION to change concentration of NO gas flowing through NOx analyzer. Then move cursor up to NOX PRECISION and press ENTER. [ESC datlogger display will now show that NO2 TITRATION electrical circuit is “closed” while the NOX PRECISION circuit is “open”.] Record time shown on ESC datalogger in logbook.

- Wait for titration (challenge) concentration values shown on ML 9841B NOx analyzer to stabilize at or near expected values. Record stable challenge concentration values for NO, NOx, and NO2 shown on “real time display w/ flags” screen of ESC datalogger in challenge blanks on level I check form on laptop. [NOTE: ESC datalogger shows these concentrations in ppm, but these concentrations must be entered in ppb in level I check form on laptop.]

- Perform second half of titration phase of NOx level I check. Use keypad on Ecotech GasCal 1000 calibrator to send 0.080 ppm of ozone (check concentration of O3 on level I check form) through NOx analyzer to complete titration phase of level I check as follows:

  - Press SELECT. Scroll down to SETUP and press SELECT. Press SELECT again to go into POINTS menu.

  - Scroll down to PT 4 and press SELECT. Scroll down to P4 O3 CONC and press SELECT.

  - Use arrow keys to change ozone concentration from 0.000 ppm to 0.080 ppm. Use SELECT to move to next position. Press ACCEPT to add 0.080 ppm of O3 to ML 9841B NOx analyzer. Press EXIT to return to main screen.

  - At main screen, use arrow keys to scroll down to O3, and make sure that screen shows proper concentration of ozone following through ML 9841B NOx analyzer.

- Wait for titration concentration values shown on ML 9841B NOx analyzer to stabilize at or near expected values shown for ozone titration on level I check form on laptop. Record stable ozone titration concentration values for NO, NOx, and NO2 shown on “real time display w/ flags” screen of ESC datalogger in challenge blanks on level I check form on laptop. [NOTE: ESC datalogger shows these concentrations in ppm, but these concentrations must be entered in ppb in level I check form on laptop.]

- Using GasCal 1000 calibrator, use arrow key to scroll down from main screen to “flow” and “source flow” values. Record flow value on “total flow” line in level I check form on laptop. Record source flow in level I check form on laptop. Then, using tag on tri-blend gas cylinder, subtract NO concentration shown in ppm on tag from total oxides of nitrogen shown in ppm on tag to get “NO2 impurity” concentration in ppm. Record this NO2 impurity value in level I check form on laptop.

- Stop flow of ozone through NOx analyzer by repeating steps 16A through 16D above, changing concentration of ozone from 0.080 ppm back to 0.000 ppm in step 16C.

- Stop titration phase of NOx level I check.  Using ESC datalogger, navigate from “real time display w/ flags” screen to “display digital outputs” screen and select ZERO.  This will begin purging of the CO, NOx, and SO2 analyzers, using zero air to flush ozone and tri-blend gas out of these analyzers. Then move cursor up to TITRATION to stop flow of tri-blend gas through the NOx analyzer. Then move cursor up and select both CO VALVE and SO2 VALVE, pressing ENTER to open each of these valves. [ESC datlogger display will now show that TITRATION electrical circuit is “open” (open electrical circuit stops flow of tri-blend gas to analyzers) while the ZERO, CO VALVE, NO VALVE, and SO2 VALVE circuits is “closed” (closed electrical circuits initiate flow of zero air through analyzers).] Record time shown on ESC datalogger in logbook.

- Replace sample inlet filters, if necessary. The sample inlet filter of NOx analyzer should be replaced after the first level I check of each month. The sample inlet filter may be located either at the back of the analyzer or inside the front cover of the analyzer. If sample inlet filter is at back of analyzer, replace it as described in steps A through C below. If sample inlet filter is front cover of analyzer, replace it as described in steps D through G below.

  - Using green plastic wrenches (which fit ribs on outside of inlet filter housing), place one wrench on orange or black half of housing and another wrench on clear half of housing. Twist wrench on orange or black half of housing counterclockwise to loosen housing. Unscrew the two halves of the housing and take them apart.

  - Remove old teflon filter from housing. Using tweezers, place new teflon filter (Savillex #1141) onto plastic screen in orange or black half of inlet filter housing. Carefully seat inner part of clear half of filter housing onto filter, making sure filter does not wrinkle or tear.

  - Screw two halves of filter housing back together. Tighten with green plastic wrenches until snug, but do not force.

  - If sample inlet filter is inside front cover of analyzer, open front cover of analyzer by pulling outward and downward on it from the top.

  - Unscrew metal cap on front of inlet filter housing and pull cap and plunger away from filter.

  - Remove rubber o-ring and old teflon filter from filter housing. Using tweezers, insert new teflon filter (Savillex #1141) into filter housing. Make sure filter isn’t wrinkled or torn. Insert rubber o-ring on top of new filter in housing.

  - Securely seat plunger onto o-ring. Screw on metal cap until hand tight. Close front cover of analyzer.

      - Restore flow of ambient air to analyzers. Wait for concentrations of NO, NOx, and NO2 shown on ML 9841B NOx analyzer to stabilize at or near zero. Then, using ESC datalogger, navigate from “real time display w/ flags” screen to “display digital outputs” screen and select NOX VALVE to stop flow of zero air through analyzers. Then select ZERO to resume flow of ambient air through the analyzers.

      - Wait for concentrations of NO, NOx, and NO2, shown on ML 9841B NOx analyzer to stabilize after flow of ambient air through system has resumed. Perform the PM10 (steps 23 through 28 below) and PM2.5 (step 29) level I check while waiting for the concentrations to stabilize.

      - Flow check of PM10 analyzer. Using ESC datalogger, navigate through “configuration menu”, “configure (data) channels”, and “put channel in maintenance” screens to put PM10 in maintenance (M) mode. Record the time this was done in logbook. Use the time shown on the ESC datalogger. Then go up on roof of air monitoring station shelter to perform flow check on inlet tube of PM10 analyzer.

      - On roof, take BGI deltaCal volumetric air flow calibrator out of its carrying case, turn it on, and let it equilibrate for a few minutes under ambient conditions.

      - Remove inlet cover (to which glass bottle is attached) from PM10 inlet. No tools should be needed to do this, but inlet cover may be difficult to remove, as it fits tightly onto the inlet tube. Check inlet cover and inlet tube for any signs of damage or debris. Remove cobwebs or other debris.

      - Attach deltaCal inlet cover to PM10 inlet. Wait for “Q value” (flow rate in liters per minute) shown on deltaCal to stabilize, then record this Q value and the time on a piece of paper.

      - Turn off deltaCal, remove its inlet cover from PM10 inlet, and put deltaCal back in its carrying case. Replace PM10 inlet cover on inlet, making sure it is pressed down as far as it will go on inlet.

      - Return to interior of air monitoring station shelter. Record Q value and time in level I check form on laptop.

      - Flow check of PM2.5 analyzer. Use down arrow key on main screen of TEOM control unit to find values for “main flow” and “auxiliary flow”. Add these two values together (sum should be close to 16.67) and record the sum in level I check form on laptop along with the time from the ESC datalogger.

      - Replace microbalance filter in TEOM sensor, if necessary. Main screen on TEOM control unit shows percent blockage of microbalance filter in TEOM sensor unit. If blockage is 85% or higher, and/or if red “check status” light is on, then check the status screen on TEOM control unit to see any messages about filter status (for example, “filter exchange needed”). Follow steps 19 through 25 below to exchange filter.

      - Using ESC datalogger, navigate through “configuration menu”, “configure (data) channels”, and “put channel in maintenance” screens to put PM2.5 into maintenance (M) mode. Record the time this was done in logbook. Use the time shown on the ESC datalogger.

      - Press “data stop” on TEOM control unit. Open front cover of TEOM sensor unit. Then open microbalance box inside sensor unit by releasing latch on top of box and pulling downward. This exposes microbalance filters inside box. The active filter is attached to the center flow inlet, while one or two backup filters are attached to standby pegs beside the center flow inlet.

      - Use microbalance filter fork to remove dirty microbalance filter from center flow inlet. Do not twist or force filter off center flow inlet. Use same fork to transfer one of the backup filters from a standby peg to the center flow inlet. Make sure fork is clean. Be careful not to damage filter or filter neck when attaching filter to center flow inlet. Press down firmly on center of filter (not at an angle) to ensure filter is seated properly.

      - If there is a second backup filter in the microbalance box, use the fork to transfer it from its present standby peg to the now vacant standby peg from which the new center flow inlet filter was just removed. Using the fork, obtain a new TX40 molded microbalance filter and place it on the vacant standby peg inside the microbalance box. The backup filters must acclimatize inside the microbalance box until each is ready for use on the center flow inlet.

      - Close and securely latch the microbalance box, then close the front cover of the TEOM sensor unit.

      - Check main screen of TEOM control unit. Percent balance of new microbalance filter should be well below 85% (about 20%). Red “check status” light on control unit should go away after a few minutes. Press “data stop” then “run”. Wait for status to read “OK” and “4”. (May take 30 minutes or more.)

      - Using ESC datalogger, take PM10 out of maintenance mode using procedure shown for CO and SO2 analyzers in step 26 below. Record time this was done in logbook, using time shown on ESC datalogger.

      - Take channels out of maintenance mode. Using ESC datalogger, navigate through “configuration menu”, “configure (data) channels”, and “take channel out of maintenance” screens to take NO, NOx, NO2, PM10, and PM2.5 out of maintenance (M) mode. Record the time this was done in logbook, and record this time as the “stop time” for the level I check of the NOx analyzer in the level I check form on laptop. Use the time shown on the ESC datalogger.

Level I Check Procedure for Ozone (O3)

The level I check procedure for the ozone analyzer does not use the cylinder of NO gas. Nevertheless, the level I check of the ozone analyzer can be performed simultaneously with the level I check of the NOx analyzer.

- Preliminary checks. Using arrow keys on Teledyne Instruments 400E ozone analyzer, scroll through instrument parameters to obtain slope (gain) value. Record this gain value on level I check form on laptop.

- Using arrow keys on API 400E ozone analyzer, scroll through instrument parameters to check the values for pressure, sample flow, sample temperature, photo lamp temperature, and offset. Compare the values to the nominal range values shown on reference sheet attached to end of this SOP. If any of the values exceed the nominal range values, record the value shown on the instrument in the bound logbook on site. Instrument values outside the nominal range may indicate problems with the analyzer that need to be corrected.

- Set up zero air generator and ozone generator. Plug power cord of Perma Pure zero air generator into wall outlet and turn zero air generator on. Open rotometer valve all the way. Allow zero air generator ~10 minutes to purge scrubber. If using the Thermo 165C transfer standard, use steps A through E. If using the Thermo 49C transfer standard, use steps F through J. If using the Environics 6103 transfer standard, use steps K through O.

- Plug power cord of Thermo 165C ozone generator into the back of the unit and wall outlet.

- Connect one end of the zero air tube with pressure gauge (end without pressure gauge) to the zero air generator.

- Remove all of the caps from the back of the Thermo 165C. Connect one end of the vent tube to the “vent” port and pass the free (open) end of the tube through the exhaust outlet at the base of the back shelter wall.

- Close the rotometer valve almost completely off on the zero air generator. Connect the opposite end of the zero air tube (end with pressure gauge) to the “air in” port on the back of the Thermo 165C and then power it on. Adjust the rotometer valve slowly until the pressure gauge is reading between 10 to 15 psi. Also, check the pressure on the front of the Thermo 165C, it should be 6.0 psi. If not, the pressure can be adjusted using the pressure control knob. Check both of these pressures occasionally to make sure they don’t drift over time.

- Connect one end of the sample tube to the “ozone” port on the Thermo 165C. Wait till the red heater lamp on the Thermo 165C is lit before proceeding to step 4 below. It may take up to 15 minutes for the heater lamp to come on. Once the heater lamp is on, place the ozone channel into maintenance using step 3.

- Plug power cord of Thermo 49C ozone calibrator into the back of the unit and wall outlet.

- Connect one end of the zero air tube with pressure gauge (end without pressure gauge) to the zero air generator.

- Remove all of the caps from the back of the Thermo 49C. Connect one end of the vent tube to the “vent port” and pass the free (open) end of the tube through the exhaust outlet at the base of the back shelter wall. Also, connect one end of exhaust tube to the “exhaust” port and pass the free end of tube through the exhaust outlet.

- Close the rotometer valve almost completely off on the zero air generator. Connect the opposite end of the zero air tube (end with pressure gauge) to the “zero air” port on the back of the Thermo 49C and then power it on. Adjust the rotometer valve slowly until the pressure gauge is reading between 10 to 15 psi. Check the pressure occasionally to make sure it doesn’t drift over time. Make sure internal pump for Thermo 49C comes on. To check, place finger over “exhaust” port and feel if air is coming out. If not, remove cover and spin start pump fan (pump should start working). Reattach cover.

- Connect one end of sample tube to the “ozone port” on the Thermo 49C. Wait for the ozone transfer standard to warm up. The main screen will show “ALARM” in the lower right hand corner. It may take up to 15-20 minutes for all of the alarms to go away. Once this happens, place the ozone channel into maintenance using step 3.

- Plug power cord of Environics 6103 ozone transfer standard into the back of the unit and wall outlet. Then power the unit on.

- Connect one end of the zero air tube to the zero air generator.

- Remove caps from the “reference” and “port 1” cross fitting and from the “output manifold ports.” Leave caps on “port 2” through “port 5” and on “sample in” port. Obtain the large vent hose with 3 small vent tubes. Connect two of the small vent tubes to the “output manifold ports” and the other small vent tube to the “vent port” and pass the large vent hose end through the exhaust outlet at the base of the back shelter wall. Also, connect one end of exhaust tube to the “exhaust” port and pass the free end of tube through the exhaust outlet.

- Press “conc” to enter concentration menu and scroll down to flow enter 5000 using keypad. Press “start”. “Reference” and “port 1” are joined together using a 4-way fitting (cross fitting). Connect a pressure gauge to one of the open inlets on the cross fitting and connect the other end of the zero air tube to the other open inlet. Adjust the rotometer valve slowly until the pressure gauge is reading between 10 to 15 psi. Check the pressure occasionally to make sure it doesn’t drift over time.

- Connect one end of the sample tube to the last available outlet on the “output manifold ports”. Wait for the Environics 6103 to warm up. It may take up to 15-20 minutes before the ozone transfer standard is ready to be used. Once the Envrionics 6103 is ready, place the ozone channel into maintenance using step 3.

- Put channel in maintenance mode. Using ESC datalogger, navigate through “configuration menu”, “configure (data) channels”, and “put channel in maintenance” screens to place ozone in maintenance (M) mode. Record time (shown on ESC datalogger) in logbook and in level I check form on laptop.

- Disconnect ambient air supply tube (which originates at the glass manifold) from inlet port on back of API 400E ozone analyzer. If a sample inlet filter is present on back of the ozone analyzer, then disconnect ambient air supply tube from inlet end of sample inlet filter. Connect opposite end of the sample tube from the ozone transfer standard to inlet port on back of ozone analyzer or to inlet end of sample inlet filter on back of ozone analyzer.

- Begin zero point acquisition of level I check. Record “ozone to zero” start time (as shown on ESC datalogger) in logbook and in level I check form on laptop. Wait for ozone concentration values shown on API 400E to stabilize at or near zero.

- While waiting for ozone concentration to stabilize, obtain the instrument “gain” (span or coefficient) value for ozone from main screen on API 400E instrument. Record this gain value in level I check form on laptop. Also, check to make sure the correct ozone transfer standard is listed as well as the corresponding expected ozone outputs on the level I check form.

- Record stable zero concentration for O3 shown on “real time display w/ flags” screen on ESC datalogger in zero audit blank on level I check form on laptop. Record expected ozone concentration shown on ozone transfer standard on level I check form on laptop. [Concentration has stabilized when values for instantaneous ozone concentration and time-weighted ozone concentration (average ozone concentration over time interval), as shown on API 400E, are the same.]

- Begin “span” phase of level I check for ozone.

- Thermo 165C - Turn ozone generator lamp switch on to begin generation of ozone. [NOTE: Wait until red light on ozone generator lamp switch has been on at least ten minutes before proceeding to step 9 below. Red heater lamp must also be lit before proceeding.] Settings on ozone generator should be as follows: switch set to “regular” instead of “ambient”; ozone level switch set to “A” instead of “B”; potential setting on dial should be set and locked to value indicated on calibration reference sheet taped to top of ozone generator; pressure indicator should still be showing value of 6.0. Record time (shown on ESC datalogger) in logbook.

- Thermo 49C – Press RUN till the “SET TO” line shows 400 ppb (LEVEL 1), and then press ENTER. Record time (shown on ESC datalogger) in logbook.

- Environics 6103 – Move cursor to O3 using arrow keys. Press EDIT and enter 0.400 ppm using keypad, and then press UPDATE. Record time (shown on ESC datalogger) in logbook.

- Wait for span concentration values on API 400E to stabilize at or near the expected value shown on level I check form on laptop. Record stable span concentration of O3 shown on “real time display w/ flags” screen on ESC datalogger in span audit blank on level I check form on laptop. Record expected ozone concentration shown on ozone transfer standard on level I check form on laptop.

- Begin “precision” phase of level I check of ozone.

- Thermo 165C - Set ozone level switch on ozone generator to “B” instead of “A”. Record time (shown on ESC datalogger) in logbook.

- Thermo 49C - Press RUN till the “SET TO” line shows 90 ppb (LEVEL 2), and then press ENTER. Record time (shown on ESC datalogger) in logbook.

- Environics 6103 – Move cursor to O3 using arrow keys. Press EDIT and enter 0.090 ppm using keypad, and then press UPDATE. Record time (shown on ESC datalogger) in logbook.

- Wait for precision concentration values shown on API 400E to stabilize at or near the expected value shown on level I check form on laptop. Record stable precision concentration of O3 shown on “real time display w/ flags” screen on ESC datalogger in precision audit blank on level I check form on laptop. Record expected ozone concentration shown on ozone transfer standard on level I check form on laptop.

- Purge analyzer with pure air.

- Thermo 165C - Stop ozone generation by turning lamp switch on ozone generator to “off” position. Return ozone level switch on ozone generator to “A” position. Record time (shown on ESC datalogger) in logbook. Allow zero air to pass through API 400E until concentration has stabilized at or near zero to purge the system.

- Thermo 49C – Press RUN till the “SET TO” line shows 0 ppb (ZERO), and then press ENTER. Record time (shown on ESC datalogger) in logbook. Allow zero to pass through API 400E until concentration has stabilized at or near zero to purge the system.

- Environics 6103 – Move cursor to O3 using arrow keys. Press EDIT and enter 0.000 ppm using keypad, and then press UPDATE. Record time (shown on ESC datalogger) in logbook. Allow zero to pass through API 400E until concentration has stabilized at or near zero to purge the system.

- Replace sample inlet filter, if necessary. The sample inlet filters of the ozone analyzer should be replaced at least once each month. The sample inlet filter on the 400E ozone analyzer is located inside the front cover of the analyzer. Replace the filter as described in steps A through F below.

- Open front cover of analyzer by pulling outward on black plastic locking pins at top of cover.

- Unscrew the metal ring on the filter housing. Remove the metal ring and the underlying glass plate and rubber o-ring. Remove the white teflon ring from the underlying filter.

- Remove old teflon filter. Using tweezers place new teflon filter (Savillex #1141) into filter housing. Make sure filter isn’t wrinkled or torn.

- Place white teflon on top of new filter. Notches on teflon ring must be showing (facing up, away from the filter) to allow sample air to flow through filter properly.

- Place rubber o-ring back in filter housing, and then lay glass plate on top of o-ring. Screw outer metal ring back onto filter housing (over glass plate) tightly, but do not force.

- Close analyzer cover and push black plastic locking pins back into place.

- When the ozone concentration (shown on the API 400E) has stabilized at or near zero, disconnect the sample tube of the ozone generator from the inlet port or sample inlet filter on the back of the ozone analyzer. Connect the ambient air supply tube (originating from the glass manifold) to inlet port or sample inlet filter on back of ozone analyzer.

- Turn off the ozone generator and the zero air generator once sample tubing has been disconnected for the API 400E.

- Take channel out of maintenance mode. Wait for concentration of O3 shown on API 400E to stabilize after flow of ambient air through system has resumed. Using ESC datalogger, navigate through “configuration menu”, “configure (data) channels”, and “take channel out of maintenance” screens to take ozone out of maintenance (M) mode. Record the time this was done in logbook, and record this time as the “stop time” for the level I check of the ozone analyzer in level I check form on laptop. Use time shown on ESC datalogger.

- Disconnect tubes and power cords for ozone generator and zero air generator. Pack up ozone generator and zero air generator in their carrying cases.

Completion of Level I Check Procedures

and

PM10/PM2.5 Flow Check Procedures

for

Roland Air Monitoring Station

After completion of level I check procedures for NOx and O3, and after completion of flow check procedures for PM10 and PM2.5:

- Save level I check file on laptop to Roland memory stick.

- Using printer 6200 at Roland air monitoring station, print two copies of level I check report (one for CNEP office file, one for on-site file at Roland air monitoring station).

- Vacuum site. Record departure time in logbook.